ELAN Interface
Description
        11/08
C79000-B5276-C176-04

---
Contents
1
General ........................................................................ 1-1
1.1
Features ........................................................................ 1-2
1.2
Broadcast Operation..................................................... 1-2
1.3
Interface Parameters..................................................... 1-3
2
ELAN Network........................................................... 2-4
2.1
Principle Connection.................................................... 2-4
2.2
Elements of ELAN Network ........................................ 2-5
2.3
Structure of ELAN Network ........................................ 2-5
2.4
Order Numbers............................................................. 2-6
3
Protocol ....................................................................... 3-1
3.1
Protocol Steps............................................................... 3-1
3.2
Features ........................................................................ 3-1
3.3
CRC-16 Checksum (Cyclic Redundancy Check) ........ 3-4
4
Commands .................................................................. 4-1
4.1
General ......................................................................... 4-1
4.2
List of all Commands ................................................... 4-2
4.3
Data Formats ................................................................ 4-8
4.4
Set Channel State ......................................................... 4-9
4.5
Set Switching Function .............................................. 4-14
4.6
Read/ Set Control Function........................................ 4-20
4.7
Read/ Write Values .................................................... 4-38
4.8
Read Diagnostic Values ............................................. 4-54
4.9
Other Commands........................................................ 4-57
5
Channel Modes ......................................................... 5-66
5.1
General ....................................................................... 5-66
5.2
Graphic Overview ...................................................... 5-66
6
Examples ..................................................................... 6-1
7
Appendix ..................................................................... 7-4
7.1
Error numbers for the command
'read logbook entry' ('k',10) .......................................... 7-4
7.2
Boot Program ............................................................... 7-8
ELAN Interface Description
II

---
ELAN Network
1
General
This ELAN interface description (Economical Local Area
Network) is valid for the SIEMENS analyzers:
•
ULTRAMAT 6 / 23
•
OXYMAT 6 / 61 / 64
•
CALOMAT 6 / 6 TG / 62
•
FIDAMAT 6
OXYMAT 61 behaves like OXYMAT 6. For information see
OXYMAT 6.
Series 6 consists of all mentioned analyzers except ULTRAMAT
23.
The following definitions are used in this description:
•
Analyzer: independent analyzer device with electronics,
sample chamber(s) and tubing in a housing
•
Channel: analyzer device with electronics (1 motherboard),
connected component(s) and tubing
•
Component: device comprising sample chamber or sensor
and associated electronics
The characters used in this description have the following
formats:
•
xxH: hex format
•
'x': ASCII format
•
Other characters are decimal
•
Commas are used to separate characters
Example: 54H = 'T' = 84
1-1

---
ELAN Network
1.1
Features
ELAN is designed as an economic serial interface for
transmitting measured values between analyzers (for correction
of cross-interference) and for simple PC communication for test
and service purposes.
A small network can be implemented if the requirements for
speed (data refresh rate 500 ms) and number of analyzers (max.
12) are not too high.
Communication is based on the following specifications:
•
Serial data communication (RS 485) with protocol (see
Chapter 3)
•
Bus capability: connection to up to 2 control systems/PCs and
up to 12 analyzers (the number of analyzers and components
may differ because one analyzer may measure several
components)
•
All analyzers connected to the ELAN have equal rights
•
To avoid bus conflicts, each analyzer must check the state of
the bus and stop the transmission immediately if necessary
(CSMA/CD)
•
A new command may only be sent if the previous command
has been answered completely (except 'broadcast', see Section
1.2).
1.2
Broadcast Operation
Analyzers are mainly restricted to answering requests. An
exception is the automatic cyclic transmission of all measured
values present in a channel (every 500 ms, identical to the answer
to command 'k', 2) (see Section 4.9).
The broadcast address is used as the target address (see Sections
3.2 and 4.9). All received messages with this address are neither
confirmed nor answered.
This procedure allows correction of cross-interference between
analyzers. Another feature is the reduced load of the bus as no
request telegrams are needed.
The number of components in broadcast mode is limited to 12.
The broadcast function can be switched off remotely by a
command or directly on the analyzer panel. This master/slave
operation requires that the control system/PC must take over the
correction of cross-interference.
1-2

---
ELAN Network
1.3
Interface Parameters
Level
RS485
Baud rate
9600
Data bits
8
Stop bit
1
Start bit
1
Parity
none
no ECHO
Pin assignments
XMT +
3
ULTRAMAT 6,
XMT -
8
OXYMAT 6,
switchable
OXYMAT 61,
termination
analyzer
connector / cable
OXYMAT 64,
by shorting
CALOMAT 6,
5V
shorted
CALOMAT 6 TG,
Ω
1k
7
CALOMAT 62,
Ω
221
9
FIDAMAT 6,
Ω
1k
0V
5
ULTRAMAT 23
8
XMT
3
XMT
1-3

---
ELAN Network
2
ELAN Network
2.1
Principle Connection
1
2
RS232
↔
RS485
Converter
RS232
RS485
8
5
RS485
9
Repeater
RS485
RS485
7
3
4
6
4
Fig. 22-1
Principle connection
2-4

---
ELAN Network
2.2
Elements of ELAN Network
Item
Name
1
Computer
↔
2
RS485
 RS232 converter with connection cable RS232 and ELAN [1];
↔
RS485
 USB converter with connection cable USB and ELAN [1];
↔
RS485
 Ethernet converter with connection cable Ethernet and ELAN [1];
3
ELAN network (RS485)
4
ELAN cable [2]
5
RS485 bus plug [3] with bridge
6
RS485 bus plug[3]
7
9-contact DSUB socket[4]
8
Analyzers
9
Option: RS485 repeater [5]
2.3
Structure of ELAN Network
•
Specifications of ELAN cable (item 4)
Cable impedance
100 to 300 Ohm, at frequency > 100 kHz
Cable capacitance
Typ. < 60 pF per meter
2
Wire diameter
> 0.22 mm
, corresponding to AWG 23
Type of cable
Twisted pair, 2 wires
Attenuation
Max. 9 dB along the whole cable
Shield
Copper braiding or braided shield and foil
•
The cable is easy to install and is highly resistant to noise due
to its double shield. It can be used for distances up to 500 m
without repeaters.
•
The double shield means that the bus cable is particularly
suitable for routing in electromagnetically loaded industrial
environments.
Further information:
−
EIA-485
−
ISO 8482: 1997 (Twisted Pair Multipoint
Interconnections)
−
DIN 66259
•
The first plug on an analyzer (item 5) needs a bus termination.
The termination is made using the bridges described in Section
1.3.
2-5

---
ELAN Network
•
The switches on the plugs (items 5 and 6) have to be OFF. The
internal resistances on the plugs are not needed.
Note
A repeater (item 9) should be used at the analyzer end with a
cable length of more than 500 m or high interferences.
2.4
Order Numbers
Item
Purveyor
Brief description
Type
Order No.
[1] Converter
SIEMENS AG
converter :
(item 2)
↔
C79451-Z1589-U1
RS485
 RS232
A5E00852382
↔
RS485
 USB
A5E00852383
↔
RS485
 Ethernet
[2] Cable
SIEMENS AG
Bus cable for
Bus cable
6XV1 830-0AH10
(item 4)
PROFIBUS
[3] Plug
SIEMENS AG
SIMATIC
6ES7972-0BB20-
(items 5, 6)
bus connector
0XA0
[4] 9 pin DSUB
Commercially available
socket (item 7)
[5] Repeater
Wieseman & Theis GmbH
RS422 isolator/
#40 10344 66201 3
(item 9)
Wittener Str. 312
RS485 repeater
D-42279 Wuppertal
1 kV isolated
(Germany)
Type 66201
2-6

---
Protocol
3
Protocol
3.1
Protocol Steps
Source
Target
Comments
DLE (10H),
Start
SOH (01H)
USED DATA
Max. characters: 68;
(target address,
10H is doubled every time
source address,
collective state,
(only when channel answers)
channel state,
(only when channel answers)
command,
data)
DLE (10H),
End of transmission
ETX (03H)
BCC, BCC
CRC-16 checksum of all transmitted
characters from DLE + SOH onwards
DLE (10H),
Confirms communication
ACK (06H) or
NAK (15H)
NAK with checksum error
The confirmation of the communication does not occur if the
broadcast address is used as the target address.
3.2
Features
The block timeout is 500 ms. The block timeout is the time in
Timeout
which the answer must have been started.
The confirm timeout is 50 ms. The confirm timeout is the time in
which the confirm communication (DLE, ACK/NAK) must have
been started.
The character timeout is 5 ms. The character timeout is the
maximum time from character to character within the string from
start (DLE, SOH) to checksum (BCC, BCC).
The maximum useful data length is 68 characters, exceeding data
Block length
will be ignored.
Each control character is started with DLE (10H). A 10H within
Control
the useful data will be doubled (10H, 10H). The transmission is
character
code-transparent as a result of this.
General
3-1

---
Protocol
Every character outside of the defined string from start (DLE,
SOH) to checksum (BCC, BCC) and outside of the confirm
communication (DLE, ACK/NAK) are ignored.
The address consists of a channel address (device address with
Addressing
controllers) and a component address:
channel address
* 16 + component address
= address (1 byte)
  | |
0 - 8
(components 1 - 9;
  | |________
depending on channel)
  |
(only checked if
  |
channels are used)
  |
1 - 12
(channels 1 - 12)
  |_________________________
13
(DCS/PC)
14
(service PC)
15
(broadcast address)
Example
ULTRAMAT 23 with HC, CO
 , NO, CO
ULTRAMAT 6 with CO, CO
2
2
(addresses: 10H, 11H, 12H, 13H)
(addresses: 20H, 21H)
ELAN
Service PC
"Dual" unit ULTRAMAT 6/OXYMAT 6 mit CO, CO
 , O
2
2
(address: E0H)
(addresses: 30H, 31H, 40H)
To find the addresses of the components from a channel, use
command 'k', 2
(= read all measured values of this channel).
The channels only react if their target address is received
correctly. If 'broadcast' is received as the target address there will
be neither confirmation nor answer.
When the channel answers, the collective state (1 byte) is
Collective state
transmitted after the addresses.
of the channel
If the collective state is 0, the transmitted measured values are
valid.
Table 33-1
Collective state of a channel
Bit 0: 1 =>
Error
Bit 1: 1 =>
Maintenance request
Bit 2: 1 =>
Not ready (not measure)
Bit 3: 1 =>
Maintenance switch on
Bit 4: 1 =>
Function check on
Bit 5: 1 =>
Command not accepted
Bit 6: 1 = >
Limit alarm
Bit 7: 1 = >
0
3-2

---
Protocol
When the channel answers, the channel state (1 byte) is
Channel state
transmitted after the collective state.
Table 3-2
Channel state
1
Warm-up
2
Pause
3
Standby
4
Measure
5
Zero calibration
6
Adjust component slope
7
Not yet defined
8
Adjust curve dip
9
Adjust linearization sensitivity
10
Adjust temperature compensation
11
Adjust pressure compensation
12
Adjust linearization zero
13
Not yet defined
14
Autocal
15
Adjust phase
16
Zero calibration of O
 sensor
2
17
Synchronous zero calibration
18
Purging for synchronous zero calibration
19
Adjust analog output
20
Adjust analog input
21
Autocal check
3-3

---
Protocol
3.3
CRC-16 Checksum (Cyclic Redundancy Check)
The CRC-16 is used for the checksum.
The block check is carried out for all transmitted characters
including control characters and DLE doubling.
The CRC-16 is performed as follows:
•
The characters to be transmitted are treated as a binary number
X.
16
•
X is multiplied by 2
 (shifted to the left sixteen times) and
16
15
2
then divided by the polynomial 2
 + 2
 + 2
 + 1. The 16-bit
remainder of this operation is the CRC-16 value.
•
This reminder is preset to FFFFH to prevent a telegram from
only consisting of zeros.
This could look as follows:
1.
Preset remainder (K0) to FFFFH
2.
XOR K0 with the first byte of the message, result to K0
3.
Shift K0 by 1 bit to the right
4.
If step 3 shifted a bit '1' to the right:
K0 XOR 0A001H
otherwise: K0 remains unchanged
5.
Repeat steps 3 and 4 eight times
6.
XOR next byte with K0
7.
Repeat steps 3 to 6 until all bytes of the message have been
processed
8.
The CRC-16 is now stored in K0
Example program using language 'C':
static word k0;
void CrcCheck (byte bb) {
byte bitanzahl;
word flag;
k0 ^= bb;
for (bitanzahl = 8; bitanzahl > 0; bitanzahl--) {
flag = k0 & 1;
k0 = k0 >> 1;
if(flag) {
k0 ^= 0xA001;
}
}
}
3-4

---
Commands
4
Commands
4.1
General
•
Commands
All commands consist of two characters.
−
1st character: ASCII letter
2nd character: a number between 1 and 255 (01H to
FFH).
−
Commands for setting/writing states or values start
with an upper-case letter ('A' to 'Z').
−
Commands for reading start with a lower-case letter ('a'
to 'z').
•
Commands for setting states or values are only executed in
remote operation. Commands for reading are always allowed.
•
Executed commands are answered with the same command.
Answer in case of command not accepted (bit 5 in collective
state is set):
−
'??'
unknown command
−
'CE'
unknown component
−
'OF'
data input or state selection is not possible
because
channel is not set to remote
−
'BS'
data input or state selection is not currently
possible
(function being executed, or wrong mode of
operation)
−
'SE'
wrong number of data
−
'DE'
wrong data value
•
There are component-related and channel-related commands.
−
Channel-related commands are accepted with any valid
component address.
•
Only reading commands are valid while the analyzer is in
calibration mode. Exceptions are the commands to control the
calibration, the abort commands Standby ('Z', 3) and Measure
('Z', 4), as well as Reset ('Z', 1).
4-1

---
 Commands
•
Data
Transmitted data are values (ASCII format, possibly with
added dimension) and control characters (1 byte between 1
and 255). All data are separated by a separation sign (1 byte:
0).
•
Excess data are ignored.
•
If data are requested within a known command which do not
belong to this channel, the answer is finished, or a blank (' ') is
inserted if other data follow.
4.2
List of all Commands
Set channel state
U6
O6
C6
C6
F6
U23
O64
C62
TG
'Z', 1
Reset
b
b
b
b
b
b
b
b
2
Pause
b
b
b
b
b
b
b
b
3
Standby
b
b
b
b
b
b
b
b
4
Measure
b
b
b
b
b
b
b
b
5
Zero calibration
k
k
k
k
k
k
k
k
6
Adjust component slope
k
k
k
k
k
k
k
k
7
Not yet defined
-
-
-
-
-
-
-
-
8
Adjust curve dip
k
-
k
k
k
k
k
k
9
Adjust linearization sensitivity
k
-
k
k
k
k
k
k
10
Adjust temperature compensation
b
b
-
-
-
b
-
-
11
Adjust pressure compensation
b
-
b
b
b
-
-
b
12
Adjust linearization zero
k
-
k
k
k
-
k
k
13
Not yet defined
-
-
-
-
-
-
-
-
14
Autocal (once)
b
b
b
b
k
b
b
b
15
Adjust phase
k
k
-
-
-
-
-
-
16
Zero calibration of O
 sensor
-
-
-
-
-
k
-
-
2
17
Not yet defined
-
-
-
-
-
-
-
-
18
Not yet defined
-
-
-
-
-
-
-
-
19
Adjust analog output
k
k
k
k
k
k
k
k
20
Adjust analog input
b
b
b
b
b
-
b
b
21
Autocal check
b
b
b
b
b
-
b
b
22
Calibration of sensor
-
-
-
-
-
-
k
-
23
Calibration of sensor % range
-
-
-
-
-
-
k
-
U6 = ULTRAMAT 6
b = channel-related command
O6 = OXYMAT 6, OXYMAT 61
k = component-related command
C6 = CALOMAT 6
# = read-only command
C6 TG = CALOMAT 6 Turbo
- = command not implemented
F6 = FIDAMAT 6
U23 = ULTRAMAT 23
4-2

---
Commands
Set switching function
U6
O6
C6
C6
F6
U23
O64
C62
TG
'F', 1
Remote on/ off
b
b
b
b
b
b
b
b
2
Pressure switch for sample gas on/ off
b
b
-
-
-
b
b
-
3
Pressure switch for reference / purge gas
b
b
-
-
-
-
b
-
on/ off
4
Total/ single calibration
k
k
k
-
k
k
k
k
5
Maintenance switch on/ off
b
b
b
b
b
b
b
b
6
Solenoid valve for zero gas on/ off
-
-
-
-
-
b
-
-
7
Not yet defined
-
-
-
-
-
-
-
-
8
Magnetic field on/ off
-
b
-
-
-
-
-
-
9
Solenoid valve for calibration gas on/ off
-
-
-
-
-
b
-
-
10
Broadcast on/ off
b
b
b
b
b
b
b
b
11
Pump on/ off
-
-
-
-
b
b
b
-
12
Radiator voltage on/ off
b
-
-
-
-
b
-
-
13
Lock logbook on/ off
b
b
b
b
b
-
b
b
14
Sample point switching on/ off
b
b
b
b
b
-
b
b
15
Suppress negative measured values on
k
k
k
k
k
b
k
k
the analog output on/ off
16
Signal violation of calibration tolerance
k
k
k
k
k
-
k
k
17
CTRL at change
-
-
-
-
-
-
b
-
18
Zero calibration before slope calibration
-
-
-
-
-
k
-
-
19
Synchronous zero calibration with
-
-
-
-
-
k
-
-
ELAN on/ off
20
Linearization on/ off
k
k
k
k
k
-
k
k
21
Temperaturecompensation on/ off
k
k
k
k
k
k
k
k
22
Pressurecompensation on/ off
k
k
k
k
k
-
k
k
23
Heating on/ off
b
b
-
-
b
-
-
b
24
Precompensation on/ off
k
-
-
-
-
-
-
-
25
Aftercompensation of the zero point on/
k
k
k
k
k
-
k
k
off
26
Aftercompensation of the span on/ off
k
k
k
k
k
-
k
k
27
Fault/ Maint. request/function control
k
k
k
k
k
-
k
k
according to NAMUR on/ off
28
Zero calibration concerted on/ off
b
-
-
-
-
-
-
-
29
Suppress negative measured values on
b
b
b
b
b
-
b
b
the display on/ off
30
Ignition on/ off
-
-
-
-
b
-
-
-
31
Heating of catalyzer on/ off
-
-
-
-
b
-
-
-
32
Fan on/ off
-
-
-
-
b
-
-
-
33
Relay by PROFIBUS on/ off
b
b
b
b
b
b
b
b
34
Constant cross-interference
-
-
-
-
-
k
-
-
35
Temperaturecompensation module on/
-
-
-
-
-
-
-
b
off
36
Not yet defined
-
-
-
-
-
-
-
-
37
Linearization module on/ off
-
-
-
-
-
-
-
b
38
internal cross-interference 1 on/ off
-
-
-
-
-
k
-
-
39
internal cross-interference 2 on/ off
-
-
-
-
-
k
-
-
4-3

---
 Commands
Read/ set control function
U6
O6
C6
C6
F6
U23
O64
C62
TG
's'/'S', 1
Limit 1
k
k
k
k
k
k
k
k
2
Limit 2
k
k
k
k
k
k
k
k
3
Analog output range
b
b
b
b
b
b
b
b
4
Relay outputs (standard)
b
b
b
b
b
b
b
b
5
Relay outputs (with optional board)
b
b
b
b
b
b
b
b
6
Binary inputs (standard)
b
b
b
b
b
b#
b
b
7
Binary inputs (with optional board)
b
b
b
b
b
b
b
b
8
Current measuring range
k
k
k
k#
k
k
k
k
9
Not yet defined
-
-
-
-
-
-
-
-
10
Measuring head heating
-
b#
-
-
-
-
-
-
11
Autocal mode
b
b
b
-
b
-
b
b
12
Autocal steps
b
b
b
-
b
b#
b
b
13
Compensation/ calibration step
k
k
k
-
k
k
k
k
14
Not yet defined
-
-
-
-
-
-
-
-
15
Pressure compensation
b
b
b
b
-
-
-
b
16
Correction 1 of cross-interference
k
k
b
b
b
k
b
b
17
Not yet defined
-
-
-
-
-
-
-
-
18
Not yet defined
-
-
-
-
-
-
-
-
19
Save measured value
b
b
b
b
b
b
b
b
20
Valves
b
b
b
b
b
b
b
b
21
AK parameters
b
b
-
-
b
-
-
-
22
Sync input/ pump
-
-
-
-
-
b
-
-
23
External interference component 1,2
k
k
b
b
b
k
b
b
24
Dimension of measured value
k
k#
k#
-
k
k
k
k#
25
Internal interference component
-
-
-
-
-
k
-
-
26
Relay outputs/ binary inputs
b#
b#
b#
b#
b#
-
b#
b#
27
Language selection
b
b
b
b
b
b
b
b
28
Automatic temperature compensation
b
b
-
-
-
-
-
-
29
Heating parameters
b
b
-
-
-
-
-
b
30
Limit 3
k
k
k
k
k
-
k
k
31
Limit 4
k
k
k
k
k
-
k
k
32
PROFIBUS parameters
b
b
b
b
b
b
b
b
33
Integrated optional board
b#
b#
b#
b#
b#
b#
b#
b#
34
Relais outputs/ binary inputs general
b#
b#
b#
b#
b#
b#
b#
b#
35
Correction 2 of cross-interference
-
-
b
b
b
-
b
b
36
Correction 3 of cross-interference
-
-
b
b
b
-
b
b
37
Correction 4 of cross-interference
-
-
b
b
b
-
b
b
38
External interference component 3,4
-
-
b
b
b
-
b
b
39
External channel of pressure
b
b
b
b
-
-
-
b
compensation
40
PROFIBUS profile
b
b
b
b
b
b
b
b
41
Internal valves
-
-
-
-
b
-
-
-
42
Relay by PROFIBUS
b#
b#
b#
b#
b#
b#
b#
b#
4-4

---
Commands
Read/ write values
U6
O6
C6
C6
F6
U23
O64
C62
TG
'w'/'W', 1
Start-of-scale values
k
k
k
k#
k
k#
k
k
2
Full-scale values
k
k
k
k#
k
k
k
k
3
Slope gas concentrations
k
k
k
k#
k
k
k
k
4
Linearization gas concentrations
k
-
k
k
k
k
k
k
5
Zero gas concentration
k
k
k
k
k
k#
k
k
6
Autocal cycle parameters
b
b
b
-
b
b
b
b
7
Purge times for Autocal steps 1 to 6
b
b
b
-
b
b
b
b
8
Purge times for Autocal steps 7 to 12
b
b
b
-
b
-
b
b
9
Limit 1
k
k
k
k
k
k
k
k
10
Limit 2
k
k
k
k
k
k
k
k
11
Integration times
k
k
k
k
k
k
k
k
12
Autorange hysteresis (% measuring
k
k
k
-
k
k
k
k
range)
13
Pump capacity
-
-
-
-
-
b
-
-
14
Date of O
sensor installation
-
-
-
-
-
b
-
-
2
15
Not yet defined
-
-
-
-
-
-
-
-
16
Time
b
b
b
b
b
-
b
b
17
Not yet defined
-
-
-
-
-
-
-
-
18
Barometric pressure
b
b
b#
b#
b#
b
b#
b#
19
Analog output
k
k
k
k
k
k
k
k
20
LCD contrast
b
b
b
b
b
b
b
b
21
Frequency
b
b
-
-
-
b
-
-
22
Reduction value
k
k
-
-
-
k
-
-
23
Phase
k#
k
-
-
-
-
-
-
24
Noise signal suppression duration
k
k
k
k
k
-
k
k
25
Calibration tolerances
k
k
k
-
k
b
k
k
26
Shock compensation
-
b
-
-
-
-
-
-
27
Parameters of external pressure sensor
b
b
b
b
-
-
-
b
28
Parameters of external interfering gas 1
k
k
b
b
b
k
b
b
29
Sample point times
b
b
b
-
b
-
b
b
40
Autorange lower limits (absolute value)
k
k
k
-
k
-
k
k
41
Autorange upper limits (absolute value)
k
k
k
-
k
-
k
k
42
Full-scale value of linearization curve
k
k#
k
k
k
k
k
k
43
Deviation in zero and slope calibration
k#
k#
k#
k#
k#
k#
k#
k#
44
Parameters of external interfering gas 2
-
-
b
b
b
k
b
b
45
Parameters of internal interfering gas 1
-
-
-
-
-
k
-
-
46
Parameters of internal interfering gas 2
-
-
-
-
-
k
-
-
47
Temperature values of temperature
k
k
k
k
k
k
k
k
comp.
48
Zero point values of temperature comp.
k
k
k
k
k
k
k
k
49
End point values of temperature comp.
k
-
k
k
k
-
k
k
50
Zero point coefficients of temperature
k
k
k
k
k
k
k
k
comp.
51
End point coefficients of temperature
k
k
k
k
k
k
k
k
comp.
52
Switching temperatur of temperature
k
-
k
k
k
k
k
k
comp.
53
Coefficients of linearization
k
k
k
k
k
k
k
k
54
Pressure values of pressure
k
-
k
k
k
-
-
k
compensation
55
Zero point values of pressure
k
-
k
k
k
-
-
k
compensation
4-5

---
 Commands
Read/ write values
U6
O6
C6
C6
F6
U23
O64
C62
TG
56
End point values of pressure
k
-
k
k
k
-
-
k
compensation
57
Coefficients of pressure compensation
k
-
k
k
k
k
-
k
58
Reference of pressure compensation
k
-
k
k
k
-
-
k
59
Aftercompensation of the zero point
k
k
k
k
k
-
k
k
60
Aftercompensation of the measured
k
k
k
k
k
-
k
k
value
61
Precompensation
k
-
-
-
-
-
-
-
62
Parameters of calibration
k
-
k
k
k
-
k
k
63
Limit 3
k
k
k
k
k
-
k
k
64
Limit 4
k
k
k
k
k
-
k
k
65
Autocal check calibration tolerances
b
b
b
-
b
-
b
b
66
Parameters of external interfering gas 3
-
-
b
b
b
-
b
b
67
Parameters of external interfering gas 4
-
-
b
b
b
-
b
b
68
Value of interfering gas 1
k
k#
b
b
b
-
b
b
69
Value of interfering gas 2
-
-
b
b
b
-
b
b
70
Value of interfering gas 3
-
-
b
b
b
-
b
b
71
Value of interfering gas 4
-
-
b
b
b
-
b
b
72
Coefficients of external interfering gas 1
-
-
b
b
b
-
b
b
73
Coefficients of external interfering gas 2
-
-
b
b
b
-
b
b
74
Coefficients of external interfering gas 3
-
-
b
b
b
-
b
b
75
Coefficients of external interfering gas 4
-
-
b
b
b
-
b
b
76
Zero-scale value of linearization curve
k
k#
k
k
k
k
k
k
77
Noise signal suppression level
k
k
k
k
k
-
k
k
78
Offset of pressure sensors
-
-
-
-
b
-
-
-
79
Parameters of preamplifier
-
-
-
-
b
-
-
-
80
Heating parameters of FIDAMAT
-
-
-
-
b
-
-
-
81
Set values of pressures
-
-
-
-
b#
-
-
-
82
Tolerances of pressures
-
-
-
-
b
-
-
-
83
Drift values
k
k
k
k
k
k
k
k
84
Heating of sensor
-
-
-
-
-
-
b
b
85
Offset / span of sensor
-
-
-
-
-
-
b
b
86
Parameter of constant cross-interference
-
-
-
-
-
k
-
-
87
Parameters of internal interfering gas 1;
-
-
-
-
-
k
-
-
new
88
Parameters of internal interfering gas 2;
-
-
-
-
-
k
-
-
new
Read diagnostic values (see chapter 4.8)
U6
O6
C6
C6
F6
U23
O64
C62
TG
'h', 1 - n
4-6

---
Commands
Other commands
U6
O6
C6
C6
F6
U23
O64
C62
TG
'k', 1
Read measured value of one component
k
k
k
k
k
k
k
k
'k', 2
Read all measured values of the channel
b
b
b
b
b
b
b
b
'k', 3
Read channel functions
b
b
b
b
b
b
b
b
'k', 4
Read component functions
k
k
k
k
k
k
k
k
'k', 5
Read error state
b
b
b
b
b
b
b
b
'k', 6
Read channel version
b
b
b
b
b
b
b
b
 'K', 7
Reset linearization coefficients
-
-
-
-
-
k
-
-
'k', 8
Read remaining time
-
-
-
-
-
b
-
-
'k', 9
Read time to next zero calibration
-
-
-
-
-
b
-
-
'k', 10
Read logbook entry
b
b
b
b
b
b
b
b
 'K', 10
Acknowledge logbook entry
b
b
b
b
b
b
b
b
'k', 11
occupied
-
-
-
-
-
-
-
-
'k', 12
Read maintenance request state
b
b
b
b
b
b
b
b
'K', 13
Acknowledge logbook error
b
b
b
b
b
b
b
b
 'K', 14
Transfer data sectors
b
b
b
b
b
b
b
b
'k', 15
occupied
-
-
-
-
-
-
-
-
'k'/'K', 16
Channel name
b
b
b
b
b
b
b
b
'k'/'K', 17
Not yet defined
-
-
-
-
-
-
-
-
'k', 18
Read message state
b
b
b
b
b
-
b
b
'k', 19
Read current Autocal step
b
b
b
b
b
-
b
b
'k', 20
Scan boot program errors
b
b
b
b
b
b
b
b
 'K', 20
Start boot program
b
b
b
b
b
b
b
b
 'K', 21
Run firmware update
b
b
b
b
b
b
b
b
 'K', 22
Clear logbook
b
b
b
b
b
-
b
b
'k', 23
Read state of the external connections
b
b
b
b
b
-
b
b
'k'/'K',24
Download mode
b
b
b
b
b
b
b
b
'K', 25
Data transfer for download
b
b
b
b
b
b
b
b
'k', 26
Set of data read/ write
b
b
b
b
b
b
b
b
'k', 27
Transfers a part of a set of data
b
b
b
b
b
b
b
b
'k', 28
Read firmware versions
b
b
b
b
b
b
b
b
'k', 29
Read warm-up state
-
-
-
-
b
-
-
-
'k'/'K', 30
Startup state
-
-
-
-
b
-
-
-
'k', 31
Read channel variant
b
b
b
b
b
b
b
b
'k', 32
Read module identification
-
-
-
-
-
-
b
b
'k'/'K', 33
Module user data
-
-
-
-
-
-
b
b
4-7

---
 Commands
4.3
Data Formats
Transmitted data are values and control characters which are
separated by 0H.
•
ASCII value ('0' - '9', possibly sign,
Value =>
decimal point).
Any scanned value (read value)
additionally contains the dimension (1
byte).
Values to be set (write value) are
always without a dimension since the
dimensions cannot be changed.
•
1-byte value (1 - 255; no 0 to avoid
Control character
separator)
(crc.) =>
Table 4-1
Dimensions
1
No dimension (only number)
21
S/ cm
41
Hz
2
ppm
22
mS/ m
42
pH
μ
3
ppb
23
mS/ cm
43
g/ l
μ
4
vpm
24
44
mg/ l
S/ m
μ
5
ppm C
25
45
l/ min
S/ cm
1
μ
6
ppm C
26
S
46
A
3
3
7
ppm C
27
min
47
mg/ dm
6
3
8
mg C/ m
28
H
48
kPa
3
Ω
9
mg/ m
29
Pa
49
k
 * cm
Ω
10
%
30
MA
50
M
 * cm
μ
11
% vol
31
51
°
V
12
% of measuring range
32
MV
52
1/ min
13
% saturation
33
V
53
1/ m
3
°
14
34
mbar
54
g/ m
%/
C
15
%/ K
35
hPa
55
g/ l
16
% weight
36
Ml/ min
56
% Vol C
Ω
17
mV/ pH
37
k
Ω
18
mV/ mbar
38
M
19
nA/ mbar
39
S
°
20
S/ m
40
C
Examples
Write values:
-3000 V => '-3000', 0
2.84 ppm, 39.0 °C, 1.8 => '2.84', 0, '39.0', 0, '1.8', 0
Read values:
-3000 V => '-3000', 0, 33, 0
2.84 ppm, 39.0 °C, 1.8 => '2.84', 0, 2, 0, '39.0', 0, 40, 0, '1.8', 0, 1, 0
4-8

---
Commands
4.4
Set Channel State
Command
Control
'Z', number, control character (if required)
structure
command:
Answer:
'Z', number
Example: start Measure
Control
'Z', 4
command:
Answer:
'Z', 4
Command has been or is being executed
or:
Answer:
'OF'
Command cannot be executed
(e.g. channel is not in Remote mode)
Reset
'Z', 1
The channel executes a Reset (as if switched off/ on).
Starting the boot program is possible within 1 s after Reset.
This command is always accepted.
Pause
'Z', 2
The channel starts the Pause state (pump off; gas flow off, ...),
Only valid in Standby state.
Standby
'Z', 3
The channel starts the Standby state (pump off, ...)
This command also aborts running functions (adjust component
slope, adjust temperature compensation etc.).
This command is not valid on FIDAMAT 6 versions without
pump.
Measure
'Z', 4
The channel starts the Measure state.
This command also aborts running functions (adjust component
slope, adjust temperature compensation etc.).
4-9

---
 Commands
Zero calibration
'Z', 5
Zero calibration of selected component (slope calibration of O
2
sensor). Calibration starts by opening the zero gas valve (relay).
The command 'Set compensation/ calibration step' ('S', 13)
imports the new zero point.
Only valid in Standby or Measure state. On FIDAMAT 6 only
valid in Measure state.
Adjust
'Z', 6,
crc.
|
component
|______
Measuring range (1 to 4)
slope
Slope calibration of the selected measuring range of the
addressed component (except O
 sensor). Calibration starts by
2
opening the calibration gas valve (relay).
The command 'Set compensation/ calibration step' ('S', 13)
imports the new slope.
With a total calibration, the calculated slope is also assigned to
the other measuring ranges.
Only valid in Standby or Measure state. On FIDAMAT 6 only
valid in Measure state.
On C6TG is only measuring range (1 to 3).
Adjust curve dip
'Z', 8
The channel calibrates the dip for linearization of the addressed
component (except O
 sensor). The calibration gas must be
2
connected to the sample gas inlet.
The command 'Set compensation/ calibration step' ('S', 13)
imports the new curve dip value.
Only valid in Standby or Measure state. On FIDAMAT 6 only
valid in Measure state.
Adjust
'Z', 9
linearization
The channel calibrates the sensitivity for linearization of the
sensitivity
addressed component (except O
 sensor). The calibration gas
2
must be connected to the sample gas inlet.
The command 'Set compensation/ calibration step' ('S', 13)
imports the new sensitivity.
Only valid in Standby or Measure state. On FIDAMAT 6 only
valid in Measure state.
4-10

---
Commands
Adjust
'Z', 10
temperature
The channel carries out a temperature measuring cycle to
compensation
determine the temperature coefficients.
The command 'Set compensation step' ('S', 13) controls this
procedure. Any temperature step must be set for at least 5 hours.
Only valid in Standby or Measure state.
Adjust pressure
'Z', 11
compensation
The channel carries out a pressure measuring cycle to determine
the pressure coefficient.
The command 'Set compensation step' ('S', 13) controls this
procedure.
Only valid in Standby or Measure state.
Adjust
'Z', 12
linearization zero
The channel carries out a sensitivity calibration for linearization
of the addressed component (except O
 sensor). The zero gas
2
must be connected to the sample gas inlet.
The command 'Set compensation/ calibration step' ('S', 13)
imports the new zero point.
Only valid in Standby or Measure state. On FIDAMAT 6 only
valid in Measure state.
Autocal
'Z', 14
The channel starts a single Autocal.
Only valid in Standby or Measure state and with channel function
Autocal on. On FIDAMAT 6 only valid in Measure state.
Commands for Autocal are only permissible for a series 6
analyzer if the optional board is fitted.
Adjust phase
'Z', 15
The component carries out an adjustment to determine the phase.
4-11

---
 Commands
Zero calibration
'Z', 16
of O
sensor
2
The channel carries out a zero calibration for the O
 sensor. Zero
2
gas must be connected via the sample gas inlet.
The command 'Set compensation/ calibration step' ('S', 13)
imports the new zero point.
Only valid in Standby or Measure state.
Adjust analog
'Z', 19
output
The component starts the calibration of the analog current output.
The command 'Set analog output' ('W', 19) controls this
procedure.
Only valid in Standby or Measure state.
Adjust analog
'Z', 20
input
The channel starts the calibration of the analog current inputs.
The analog current output 1 must be adjusted beforehand.
Calibration steps :
1.
Connect analog current output 1 to analog current input 1.
Send command 'Z', 20.
2.
Send command 'Set compensation/ calibration step' ('S', 13)
(analog current input 1 is adjusted to 20 mA).
3.
Send command 'Set compensation/ calibration step' ('S', 13)
(analog current input 1 is adjusted to 0 mA).
4.
Connect analog current output 1 to analog current input 2.
5.
Send command 'Set compensation/ calibration step' ('S', 13)
(analog current input 2 is adjusted to 20 mA).
6.
Send command 'Set compensation/ calibration step' ('S', 13)
(analog current input 2 is adjusted to 0 mA).
Only valid in Standby or Measure state.
Autocal check
'Z', 21
The channel starts a single Autocal check.
Only valid in Standby or Measure state and with channel function
Autocal on. On FIDAMAT 6 only valid in Measure state.
Commands for Autocal check are only permissible if the optional
board is fitted.
4-12

---
Commands
Calibration of
'Z', 22,
crc.
sensor
The component starts the calibration of the sensor.
Air and the selected calibration gas are needed.
The measured value of sensor can be read by command 'h',6.
Calibration steps :
1.
Connect air. The command 'Set compensation/ calibration
step' ('S', 13) imports the new zero point of sensor.
2.
Connect selected calibration gas for the vpm range. The
command 'Set compensation/ calibration step' ('S', 13) imports
the new span of sensor for the vpm range and finishes the
calibration.
Only valid with total calibration and in Standby or Measure state.
Calibration of
'Z', 23,
crc.
sensor % range
The component starts the calibration of the sensor %Vol range.
The selected calibration gase are needed.
The measured value of sensor can be read by command 'h',6.
Connect selected calibration gas for the %Vol range. The
command 'Set compensation/ calibration step' ('S', 13) imports the
new span of sensor for the %Vol range and finishes the
calibration.
Only valid with total calibration and in Standby or Measure state.
4-13

---
 Commands
4.5
Set Switching Function
Command
Control
'F', number, function ('0' => off; '1' => on)
structure
command:
Answer:
'F', number
Example: Remote on
Control
'F', 1, '1', 0
command:
Answer:
'F', 1
Remote on/ off
'F', 1,
function
   |
   |______
'0' => off; '1' => on
The channel starts or finishes Remote mode (remote control via
the interface).
Without Remote, only read commands are accepted.
Only valid if the channel is coded (all codes switched off).
Pressure switch
'F', 2,
function
   |
for sample gas
   |______
'0' => off; '1' => on
on/ off
The channel starts or finishes monitoring the sample gas pressure.
Pressure switch
'F', 3,
function
   |
for reference /
   |______
'0' => off; '1' => on
purge gas on/ off
The channel starts or finishes monitoring the reference / purge
gas pressure.
Total/ single
'F', 4,
function
   |
'0' => single calibration;
calibration
   |______
'1' => total calibration
The component performs with total calibration (calibration valid
for all ranges) or single calibration (each range is calibrated
independent of the others).
The function is not available at O2-component of ULTRAMAT
23.
4-14

---
Commands
Maintenance
'F', 5,
function
   |
switch on/ off
   |______
'0' => off; '1' => on
The channel is being serviced (a code has been entered).
Solenoid valve
'F', 6,
function
   |
for zero gas on/
   |______
'0' => off; '1' => on
off
The channel switches the internal solenoid valve and the relay
contact for the external solenoid valve on or off.
Magnetic field
'F', 8,
function
   |
on/ off
   |______
'0' => off; '1' => on
The channel switches the magnetic field on or off.
If the magnetic field is switched off, the heating of the measuring
head is also off.
Solenoid valve
'F', 9,
function
   |
for calibration
   |______
'0' => off; '1' => on
gas on/ off
The channel switches the relay contact for the external solenoid
valve on or off.
Broadcast
'F', 10,
function
   |
on/ off
   |_____
'0' => off; '1' => on
The channel starts or finishes the automatic transmission
(broadcasting) of measured values (identical to answer to
command 'k', 2).
Pump
'F', 11,
function
   |
on/ off
   |_____
'0' => off; '1' => on
The channel switches the internal pump on or off.
4-15

---
 Commands
Radiator voltage
'F', 12,
function
   |
on/ off
   |_____
'0' => off; '1' => on
The channel switches the radiator voltage on or off.
Lock logbook
'F', 13,
function
   |
on/ off
   |_____
'0' => off; '1' => on
The channel stops or allows the input of further messages into the
logbook.
Sample point
'F', 14,
function
   |
switching on/ off
   |_____
'0' => off; '1' => on
The channel stops or allows sample point switching.
Suppress
'F', 15,
function
   |
negative
   |_____
'0' => off; '1' => on
measured values
on the analog
output on/ off
Signal violation
'F', 16,
function
   |
of calibration
   |_____
'0' => off; '1' => on
tolerance on/off
CTRL at
'F', 17,
function
   |
change
   |_____
'0' => off; '1' => on
on/ off
'Function check' is set while dimension of measured value is
changed.
Zero calibration
'F', 18,
function
   |
before slope
   |_____
'0' => off; '1' => on
calibration
on/ off
The zero is automatically calibrated before the component slope.
4-16

---
Commands
Synchronous
'F', 19,
funktion
   |
zero calibration
   |_____
'0' => off; '1' => on
with ELAN
on/ off
Broacast has to be switched on before.
Linearization
'F', 20,
funktion
   |
on/ off
   |_____
'0' => off; '1' => on
Temperature
'F', 21,
funktion
   |
compensation
   |_____
'0' => off; '1' => on
on/ off
With ULTRAMAT 23, the function is available only at O2-
component.
Pressure
'F', 22,
funktion
   |
compensation
   |_____
'0' => off; '1' => on
on/ off
Heating on/ off
'F', 23,
funktion
   |
   |_____
'0' => off; '1' => on
Only valid if heating is integrated.
Precompensa-
'F', 24,
funktion
   |
tion on/ off
   |_____
'0' => off; '1' => on
Aftercompensa-
'F', 25,
funktion
   |
tion of the zero
   |_____
'0' => off; '1' => on
point on/ off
Aftercompensa-
'F', 26,
funktion
   |
tion of the span
   |_____
'0' => off; '1' => on
on/ off
4-17

---
 Commands
Fault / Maint.
'F', 27,
funktion
   |
request /
   |_____
'0' => off; '1' => on
Function control
according to
NAMUR on/ off
Zero calibration
'F', 28,
funktion
   |
concerted on/ off
   |_____
'0' => off; '1' => on
A zero calibration performs a calibration of both components.
This command applies only to ULTRAMAT 6 with two
components.
Suppress
'F', 29,
function
   |
negative
   |_____
'0' => off; '1' => on
measured values
on the display
on/ off
Ignition
'F', 30,
function
   |
on/ off
   |_____
'0' => off; '1' => on
If flame is burning the ignition is switched off.
Heating of
'F', 31,
funktion
   |
catalyzer on/ off
   |_____
'0' => off; '1' => on
Only valid for FIDAMAT with integrated catalyzer.
Fan
'F', 32,
funktion
   |
on/ off
   |_____
'0' => off; '1' => on
Relay by
'F', 33,
funktion
   |
PROFIBUS
   |_____
'0' => off; '1' => on
on/ off
This command is only permissible with the PROFIBUS optional
board fitted and unused relay outputs of optional board.
4-18

---
Commands
Constant cross-
'F', 34,
funktion
   |
interference
   |_____
'0' => off; '1' => on
on/ off
Temperature
'F', 35,
funktion
   |
compensation
   |_____
'0' => off; '1' => on
module on/ off
Linearization
'F', 37,
funktion
   |
module on/ off
   |_____
'0' => off; '1' => on
Internal cross-
'F', 38,
funktion
   |
interference 1
   |_____
'0' => off; '1' => on
on/ off
Internal cross-
'F', 39,
funktion
   |
interference 2
   |_____
'0' => off; '1' => on
on/ off
4-19

---
 Commands
4.6
Read/ Set Control Function
Read control:
Command
structure
Control
's', number
command:
Answer:
's', number, crc.
Example: analog output range
Control
's', 3
command:
Answer:
's', 3, 81H, 0
  | | |
  | | |_
Separator
  | |____
Range: 2 - 20 mA, not inverted
  |_______
Analog output current range
Set control:
Control
'S', number, crc.
command:
Answer:
'S', number
Example: limit 1
Control
'S', 1, A3H, 0
command:
Answer:
'S', 1
Read/ set limit 1
's'/'S', 1,
crc.
 |
 |_
Bit 7: 1
Bit 6: 0
Bit 5: 1 => limit active
(0 => passive)
Bit 4: 1 => active upward violation (0 => downward violation)
Bit 3: 1 => active in meas. range 4
(0 => passive)
Bit 2: 1 => "
"" 3
"
Bit 1: 1 => "
"" 2
"
Bit 0: 1 => "
"" 1
"
The range information is ignored for the ULTRAMAT 23
(always measuring ranges 1 and 2).
The data for upward/ downward violation for the ULTRAMAT
23 are only observed if "Limit active" is also present.
There has to be at least one active range for the analyzers of
Series 6.
There is only meassure range 1 to 3 available for the C6TG
4-20

---
Commands
Read/ set limit 2
(as for alarm 1)
's'/'S', 2,
crc.
Read/ set analog
's'/'S', 3,
crc.
 |
output range
 |_
Bit 7: 1
Bit 6, 5, 4, 3 : 0
Bit 2: 1 => inverted
 (0 => not inverted)
Bit 1
Bit 0
1
1 => 4 - 20 mA Namur
1
0 => 4 - 20 mA
0
1 => 2 - 20 mA
0
0 => 0 - 20 mA
The ranges are never inverted with the ULTRAMAT 23.
With ULTRAMAT 23, the output range Namur can only be set
with Softwareversion 2.9.2 or later.
Read/ set relay
's'/'S', 4,
crc. 1, crc. .2, crc. 3, crc. 4, crc. 5, crc. 6, crc. 7, crc. 8
outputs
 | | | | | | | |
Relay 8
 | | | | | | | |____
(standard)
Relay 7
 | | | | | | |_________
Relay 6
 | | | | | |______________
Relay 5
 | | | | |___________________
Relay 4
 | | | |________________________
Relay 3
 | | |_____________________________
Relay 2
 | |__________________________________
Relay 1
 |_______________________________________
Control
Function of the relay output
Comment
character
(crc.)
1
Free
2
Fault
3
Maintenance request
4
Calibration
Only series 6
5
Component 1 measuring range 1 selected
Only series 6
6
Component 1 measuring range 2 selected
7
Component 1 measuring range 3 selected
Only series 6
8
Component 1 measuring range 4 selected
Only series 6
9
Component 1 alarm 1 triggered
10
Component 1 alarm 2 triggered
11
Function check
12
Valve for sample gas is open
Only series 6
13
Valve for zero gas is open
14
Valve for calibration gas 1 is open
15
Valve for calibration gas 2 is open
Only series 6
16
Valve for calibration gas 3 is open
Only series 6
4-21

---
 Commands
Control
Function of the relay output
Comment
character
(crc.)
17
Valve for calibration gas 4 is open
Only series 6
18
Measuring point 1
Only series 6
19
Measuring point 2
Only series 6
20
Measuring point 3
Only series 6
21
Measuring point 4
Only series 6
22
Measuring point 5
Only series 6
23
Measuring point 6
Only series 6
24
Signal from measuring point 1
Only series 6
25
Signal from measuring point 2
Only series 6
26
Signal from measuring point 3
Only series 6
27
Signal from measuring point 4
Only series 6
28
Signal from measuring point 5
Only series 6
29
Signal from measuring point 6
Only series 6
30
Signal contact (for synchronization with
Autocal)/sync signal
31
Flow of reference / purge gas
Only U6/ O6 / O64
32
Flow of sample gas
Only U6/ O6 / O64
33
Valve for zero gas 2 is open
Only series 6
34
Component 2 measuring range 1 selected
Only U6
35
Component 2 measuring range 2 selected
Only U6/ U23
36
Component 2 measuring range 3 selected
Only U6
37
Component 2 measuring range 4 selected
Only U6
38
Component 2 limit 1 triggered
Only U6/ U23
39
Component 2 limit 2 triggered
Only U6/ U23
40
Component 3 measuring range 2 selected
Only U23
41
Component 3 limit 1 triggered
Only U23
42
Component 3 limit 2 triggered
Only U23
43
Component 4 measuring range 2 selected
Only U23
44
Component 4 limit 1 triggered
Only U23
45
Component 4 limit 2 triggered
Only U23
46
CAL/MEAS
Only U23
47
Component 1 limit 3 triggered
Only series 6
48
Component 1 limit 4 triggered
Only series 6
49
Component 2 limit 3 triggered
Only U6
50
Component 2 limit 4 triggered
Only U6
51
Heating
Only U6/ O6
52
Autocal check difference
Only series 6
53
Comp. 2 Valve for zero gas is open
Only U6
54
Comp. 2 Valve for calibration gas 1 is open
Only U6
55
Comp. 2 Valve for calibration gas 2 is open
Only U6
56
Comp. 2 Valve for calibration gas 3 is open
Only U6
57
Comp. 2 Valve for calibration gas 4 is open
Only U6
58
Valve for hydrogen is open
Only F6
59
Valve for combustion air is open
Only F6
60
Internal valve 5 is open
Only F6
61
Internal valve 6 is open
Only F6
4-22

---
Commands
Control
Function of the relay output
Comment
character
(crc.)
62
Operation request
Only F6
63
Device with dimension % Vol
Only O64 (since V0.1.6)
Series 6 analyzers have only 6 relay outputs. Each function can
be set for only one relay.
Read/ set relay
's'/'S', 5,
crc. 1, crc. 2, crc. 3, crc. 4, crc. 5, crc. 6, crc. 7, crc. 8
outputs (with
 | | | | | | | |
optional board)
Relay 8
 | | | | | | | |__
Relay 7
 | | | | | | |______
Relay 6
 | | | | | |___________
Relay 5
 | | | | |________________
Relay 4
 | | | |_____________________
Relay 3
 | | |__________________________
Relay 2
 | |_______________________________
Relay 1
 |____________________________________
See "Relay outputs (standard)" for explanation of control
characters.
Read/ set binary
's'/'S', 6,
crc. 1, crc. 2, crc. 3, crc. 4, crc. 5, crc. 6
inputs (standard)
Binary input 6
 | | | | | |_
Binary input 5
 | | | | |______
Binary input 4
 | | | |___________
Binary input 3
 | | |________________
Binary input 2
 | |_____________________
Binary input 1
 |__________________________
Control
Function of the binary input
Comment
character
1
Vacant
2
Other fault (external fault)
3
Fault probe heating (external fault)
4
Fault sample gas filter (external fault)
5
Fault gas cooler (external fault)
6
Other maintenance request
(external maintenance request)
7
Maintenance request probe heating
(external maintenance request)
8
Maintenance request sample gas filter
(external maintenance request)
4-23

---
 Commands
Control
Function of the binary input
Comment
character
9
Maintenance request gas cooler
(external maintenance request)
10
Acknowledge (logbook reset)
11
Function control 1
12
Function control 2
13
Function control 3
14
Function control 4
15
Autocalibration
U23 only read
16
Component 1 Range 1
17
Component 1 Range 2
18
Component 1 Range 3
Only series 6
19
Component 1 Range 4
Only series 6
20
Component 1 Autorange
21
Zero gas
Only series 6
22
Span gas
Only series 6
23
Sample gas
Only series 6
24
Zero calibration
Only series 6
25
Component 1 Span calibration
Only series 6
26
Fault condensate vessel (external fault)
27
Fault sample pump/ flow (external fault)
28
Fault sample line (external fault)
29
Maintenance request condensate vessel
(external maintenance request)
30
Maintenance request sample pump/ flow
(external maintenance request)
31
Maintenance request sample line
(external maintenance request)
32
Pump on/ off
Only U23; only read
33
Synchron zero adjustment
Only U23; only read
34
Component 2 range 1
Only U6/ U23
35
Component 2 range 2
Only U6/ U23
36
Component 2 range 3
Only U6
37
Component 2 range 4
Only U6
38
Component 2 autorange
Only U6/ U23
39
Component 3 range 1
Only U23
40
Component 3 range 2
Only U23
41
Component 3 autorange
Only U23
42
Component 4 range 1
Only U23
43
Component 4 range 2
Only U23
44
Component 4 autorange
Only U23
45
Component 2 span calibration
Only U6
46
Autocal check
Only series 6
47
Measure state locked
Only series 6
48
Component 2 zero gas
Only U6
49
Component 2 span gas
Only U6
50
Component 2 zero calibration
Only U6
Each function can be set for only one binary input.
4-24

---
Commands
Read/ set binary
's'/'S', 7,
cr. 1, cr. 2, cr. 3, cr. 4, cr. 5, cr. 6, cr. 7, cr. 8
inputs (with
Binary input 8
 | | | | | | | |_
optional board)
Binary input 7
 | | | | | | |_____
Binary input 6
 | | | | | |__________
Binary input 5
 | | | | |_______________
Binary input 4
 | | | |____________________
Binary input 3
 | | |_________________________
Binary input 2
 | |______________________________
Binary input 1
 |___________________________________
See "Binary inputs (standard)" for explanation of control
characters.
Read/ set current
's'/'S', 8,
crc.
 |
measuring range
 |_
Bit 7: 1
Bit 6, 5, 4, 3: 0
Bit 2: 1 => autorange on
Bit 1
Bit 0
1
1 => measuring range 4
1
0 => "
3
0
1 => "
2
0
0 => "
1
The autorange function is switched off if a measuring range is set.
Ranges 3 and 4 cannot be selected for the ULTRAMAT 23.
There are only measuring range 1-3 in readonly mode available,
for the C6TG.
Read measuring
's', 10,
crc.
 |
head heating
 |_
Bit 7: 1
Bit 6, 5, 4: 0
Bit 3: 1 => measuring head heating on (0 => off)
Bit 2: 0
Bit 1
Bit 0
1
1 => 100 °C setpoint temperature
1
0 => 90 °C
"
0
1 => 80 °C
"
0
0 => 75 °C
"
4-25

---
 Commands
Read/ set
's'/'S', 11,
crc.
 |
Autocal mode
 |_
Bit 7: 1
Bit 6, 5, 4 : 0
Bit 3: 1 => Autocal check on (0 => off)
Bit 2: 1 => Autocal on (0 => off)
Bit 1: 1 => start by binary input
Bit 0: 1 => start by cyclic parameter
Commands for Autocal are only valid for the Series 6 if the
optional board is fitted.
Read/ set
's'/'S', 12,
crc. 1, crc. 2, crc. 3, ... crc. 12
Autocal steps
 | | | |
Step 12
 | | | |___
   " 3
 | | |__________
   " 2
 | |_______________
   " 1
 |____________________
Control character
Step
Comment
1
Not used
2
Zero gas 1
3
Zero gas 2
4
Calibration gas 1
5
Calibration gas 2
6
Calibration gas 3
7
Calibration gas 4
8
Sample gas purging
9
Intermediate sample gas mode
10
Signalling contact
Commands for Autocal are only valid for the Series 6 if the
optional board is fitted.
Read/ set
's'/'S', 13,
crc.
 |
compensation/
 |_
1 – 255
calibration step
With multi-step compensation/ calibration procedures, a "Set
step" command sets the current values for calculation of the
compensation and starts the next step. The transmitted control
character is ignored.
With "Read step", the control character defines the current step.
4-26

---
Commands
Read/ set
's'/'S', 15,
crc. 1, crc. 2
pressure
 | |
Bit 7: 1
 | |_
compensation
Bit 6, 5, 4, 3, 2: 0
 |
Bit 1
Bit 0
Analog input of ext. pressure
 |
 |
1
0 =>
4 - 20 mA/ NAMUR
 |
0
1 =>
2 - 20 mA
 |
0
0 =>
0 - 20 mA
 |
Bit 7: 1
 |______
Bit 6, 5, 4, 3, 2: 0
Bit 1
Bit 0
1
1 => pressure comp. with ELAN
1
0 => pressure comp. with analog current 2
0
1 => pressure comp. with int. pressure
0
0 => no pressure compensation
The settings for pressure compensation are channel-related.
However with the setting 'no pressure compensation' it is possible
to switch the single components off from compensation.
With CALOMAT 6 and CALOMAT 62 , there is no pressure
compensation with internal pressure.
4-27

---
 Commands
's'/'S', 16,
crc. 1, crc. 2
Read/ set
 | |
Bit 7: 1
correction 1 of
 | |_
Bit 6: 0
 |
cross-
Bit 5: 1 => corr. of cross-interference valid for range 4
 |
interference
Bit 4: 1 => corr. of cross-interference valid for range 3
 |
Bit 3: 1 => corr. of cross-interference valid for range 2
 |
Bit 2: 1 => corr. of cross-interference valid for range 1
 |
 |
 |
Bit 1
Bit 0
Analog input of ext. interfering gas
 |
 |
1
0 =>
4 – 20 mA / NAMUR
0
1 =>
2 – 20 mA
 |
0
0 =>
0 – 20 mA
 |
 |
Bit 7: 1
 |______
Bit 6, 5, 4, 3, 2: 0
Bit 2
Bit 1
Bit 0
1
0
1 => residual gas influence with
        analog current (analog input 2)
0 => residual gas influence with
1
0
        internal component
0
1 => residual gas influence via ELAN
1
0
0 => residual gas influence with
1
        analog current (analog input 1)
0
1 => constant residual gas interference
0
0 => no residual gas influence
0
0
With ULTRAMAT 23, the correction of cross-interference is
always valid for range 1 and 2, and residual gas influence is
possible only via ELAN.
There has to be at least one range for cross-interference with
Series 6 analyzers.
The residual gas influence with internal component is only valid
with ULTRAMAT 6 with 2 components.
The residual gas influence with analog input 2 is only valid with
CALOMAT 6 and FIDAMAT 6.
With CALOMAT 6 and FIDAMAT 6, the current range of the
analog input can only be set when residual gas influence with
analog current 1 or 2 is selected.
4-28

---
Commands
Read/ set
's'/'S', 19,
crc.
 |
measured-value
 |_
Bit 7: 1
saving
Bit 6, 3, 2: 0
Bit 5
Bit 4
with error:
1
1 =>
highest value (20 mA)
1
0 =>
lowest value (0/ 2/ 4 mA)
0
1 =>
save measured value
0
0 =>
save function off (update)
Bit 1
Bit 0
with function check:
1
1 =>
highest value (20 mA)
1
0 =>
lowest value (0/ 2/ 4 mA)
0
1 =>
save measured value
0
0 =>
save function off (update)
With Series 6 analyzers there is no difference between an error
and a function check. The input for a function check thus also
applies to an error.
Read/ set valves
's'/'S', 20,
crc.
 |_
Bit 7: 1
Bit 6: 0
Bit 5: 1 => valve for calibration gas 4 open (0 => close)
Bit 4: 1 => "
 calibration gas 3 "
"
Bit 3: 1 => "
 calibration gas 2 "
"
Bit 2: 1 => "
 calibration gas 1 "
"
Bit 1: 1 => "
 zero gas
"
"
Bit 0: 1 => "
 sample gas
"
"
Only one valve can be open.
With ULTRAMAT 23 only bit 0 - 2 are used.
4-29

---
 Commands
Read/ set AK
's'/'S', 21,
crc. 1, crc. 2, crc. 3, crc. 4, crc. 5
parameters
 | | | | |
'Don't care' character
 | | | | |_
End character
 | | | |______
Start character
 | | |___________
1
AK
 format (see table)
 | |________________
Baud rate (see table)
 |_____________________
This command is only permissible with the AK optional board
fitted.
The start character, end character and 'Don't care' character must
not be 0.
Control character 1
Baud rate
1
300
2
600
3
1200
4
2400
5
4800
6
9600
AK format
Control
No. of
Parity
No. of
No. of
character 2
data bits
stop bits
bits per sign
1
7
None
2
10
2
7
Even
1
10
3
7
Odd
1
10
4
8
None
1
10
5
7
Even
2
11
6
7
Odd
2
11
7
8
Even
1
11
8
8
Odd
1
11
9
8
None
2
11
Sync input/
's'/'S', 22,
crc.
 |
pump
 |_
Bit 7: 1
(ULTRAMAT 23)
Bit 6, 5, 4, 3: 0
Bit 2: 1 => pump on with Zero calibration
(0 => off)
Bit 1: 1 => pump on with Measure
(0 => off)
Bit 0: 1 => Sync input starts Zero calibration
 (0 => no Zero calibration; only set Cal contact)
1
 AK: Arbeitskreis der deutschen Automobilindustrie (Working
Party of the German Automotive Industry)
4-30

---
Commands
Read/ set
's'/'S', 23,
crc. 1, crc. 2
external
Component address for correction of cross-
 | |__
interference 2
 |
interference
Bit 7, 6, 5, 4: channel address
(1 - 12)
 |
component 1,2
Bit 3, 2, 1, 0: component address
 (0 - 3)
 |
Component address for correction of cross-
 |______
interference 1
Bit 7, 6, 5, 4: channel address
(1 - 12)
Bit 3, 2, 1, 0: component address
 (0 - 3)
The address for correction may not be the address of the device.
With the ULTRAMAT 6 and OXYMAT 6, only the component
address for correction of cross-interference 1 is valid.
Note
The component addresses in device menu are 1 – 4 !
Read/ set
's'/'S', 24,
crc.
dimension of
measured value
Control
Dimension
Dimension
Dimension
Dimension
character
U23
 U6,O6,C6,C62
F6
O64
1
Not used
% Vol
ppm C1
% Vol
(since V0.1.6)
2
vpm
vpm
ppm C3
vpm
3
3
%
mg/ m
ppm C6
automatical
(only for read)
change
(since V0.1.6)
3
3
4
mg/ m
ppm
mg/ m
5
ppm
mg/ l
% Vol C
(only for read)
3
6
g/ m
(since V4.4.0)
7
g/ l
(since V4.4.0)
This command is not valid for the O
 sensor of the U23.
2
With the ULTRAMAT 6 and FIDAMAT 6, no change allowed if
full-scale value of linearization curve in new dimension exeeds
99999.
No change allowed with ULTRAMAT 23 components with
dimension %.
4-31

---
 Commands
Read/ set
's'/'S', 25,
crc. 1, crc. 2
internal
 | |
Component address for correction of internal cross-
 | |__
interference
interference 2
 |
component
Bit 7: 1
 |
Bit 6, 5, 4: 0
 |
Bit 3, 2, 1, 0: component address
 (0 - 3)
 |
 |
Component address for correction of internal cross-
 |______
interference 1
Bit 7: 1
Bit 6, 5, 4: 0
Bit 3, 2, 1, 0: component address
 (0 - 3)
Read relay
's'/, 26,
crc. 1, crc. 2, crc. 3, crc. 4
inputs/
 | | | |
binary inputs 8 - 14
 | | | |__
binary outputs
 | | |
Bit 7 : 1
 | | |
Bit 0 – 6 : state of binary inputs 8 - 14
 | | |
binary inputs 1 - 7
 | | |_______
Bit 7 : 1
 | |
Bit 0 – 6 : state of binary inputs 1 - 7
 | |
 | |
relay outputs 8 - 14
 | |____________
Bit 7 : 1
 |
Bit 0 – 6 : state of relay outputs 8 - 14
 |
 |
relay outputs 1 - 7
 |_________________
Bit 7 : 1
Bit 0 – 6 : state of relay outputs 1 - 7
Relay outputs 7 – 14 and binary inputs 7 – 14 are on the optional
board.
Read/ set
's'/'S', 27,
crc.
 |
language
 |_
1 => German / ordered (second) language
selection
2 => English
3 => Spanish
4 => French
5 => Italian
Only English and the ordered language (crc. = 1) can be set.
4-32

---
Commands
Automatic
's'/'S', 28,
crc. 1, crc. 2, crc. 3
temperature
 | | |
Delay of a temperature step
 | | |___
compensation
(2 - 15 h)
 | |
 | |
Lock on time at the end of a temperature step
 | |________
(2 - 59 min)
 |
 |
Bit 7: 1
 |_____________
Bit 6, 5, 4, 3, 2, 1: 0
Bit 0: 1 => automatic sequence for the
        temperatur compensation on
        (0 => off)
The parameters of the automatic temperature compensation are
lost when device is reset.
Heating
's'/'S', 29,
crc. 1, crc. 2
parameters
 | |
°
Nominal temperature (O6 : 65 - 130
C;
 | |___
        C62 : 60 - 80 °C )
 |
 |
Bit 7: 1
 |________
Bit 6, 5, 4, 3, 2, 1: 0
Bit 0: 1 => Heating on (0 => no heating)
With ULTRAMAT 6 the nominal temperature can be read only.
Limit 3
(see limit 1)
's'/'S', 30,
crc.
(see limit 1)
's'/'S', 31,
crc.
Limit 4
PROFIBUS
's'/'S', 32,
crc.
parameter
 |
Bit 7: 1
 |________
Bit 6,5,4,3,2,1,0 => PROFIBUS address (0 - 126)
This command is only permissible with the PROFIBUS optional
board fitted.
The address can only be set when there is no cyclical
communication.
4-33

---
 Commands
Integrated
's', 33,
crc.
1 => no option board fitted
optional board
 |
2 => Autocal optional board
 |________
3 => AK optional board
4 => PROFIBUS PA optional board
5 => PROFIBUS DP optional board
This command is only permissible with the PROFIBUS optional
board fitted.
The address can only be set when there is no cyclical
communication.
Relay outputs/
's'/, 34,
crc. 1, crc. 2, crc. 3, crc. 4 , crc. 5
binary inputs
 | | | | |
general
Binary inputs 8 - 14 :
 | | | | |__
Bit 7 : 1
 | | | |
Bit 0 - 6 : state of binary inp. 8 –
 | | | |
14
 | | | |
Binary inputs 1 - 7 :
 | | | |______
Bit 7 : 1
 | | |
Bit 0 - 6 : state of binary inp. 1 – 7
 | | |
 | | |
Relay outputs 15 - 16 :
 | | |___________
Bit 7 : 1
 | |
Bit 0 - 1 : state of relay 15 - 16
 | |
 | |
Relay outputs 8 - 14 :
 | |________________
 |
Bit 7 : 1
 |
Bit 0 - 6 : state of relay 8 - 14
 |
Relay outputs 1 - 7 :
 |_____________________
Bit 7 : 1
Bit 0 - 6 : state of relay 1 - 7
The relay outputs 1 - 8 and the binary inputs 1 - 6 are on the main
board.
All further relay outputs and binary inputs are on the optional
board.
ULTRAMAT 23 has no binary inputs 4 – 6.
Series 6 analyzers have no relay output 7 - 8.
4-34

---
Commands
Read/ set
's'/'S', 35,
crc. 1, crc. 2
correction 2 of
 | |
Bit 7: 1
 | |_
cross-
Bit 6: 0
 |
interference
Bit 5: 1 => corr. of cross-interference valid for range 4
 |
Bit 4: 1 => corr. of cross-interference valid for range 3
 |
Bit 3: 1 => corr. of cross-interference valid for range 2
 |
Bit 2: 1 => corr. of cross-interference valid for range 1
 |
Bit 1
Bit 0
Analog input of ext. interfering gas
 |
 |
1
0 =>
4 – 20 mA / NAMUR
0
1 =>
2 – 20 mA
 |
0
0 =>
0 – 20 mA
 |
 |
 |______
Bit 7: 1
Bit 6, 5, 4, 3, 2: 0
Bit 2
Bit 1
Bit 0
1
0
1 => residual gas influence with
        analog current (analog input 2)
0 => residual gas influence with
1
0
        internal component
0
1 => residual gas influence via ELAN
1
0
0 => residual gas influence with
1
        analog current (analog input 1)
0
1 => constant residual gas interference
0
0 => no residual gas influence
0
0
There has to be at least one range for cross-interference.
The residual gas influence with internal component is not valid.
The current range of the analog input can only be set when
residual gas influence with analog current 1 or 2 is selected.
Read/ set
 (see correction 2)
's'/'S', 36,
crc. 1, crc. 2
correction 3 of
cross-
interference
Read/ set
(see correction 2)
's'/'S', 37,
crc. 1, crc. 2
correction 4 of
cross-
interference
4-35

---
 Commands
Read/ set
's'/'S', 38,
crc. 1, crc. 2
external
Component address for correction of cross-
 | |__
interference 4
 |
interference
Bit 7, 6, 5, 4: channel address
(1 - 12)
 |
component 3,4
Bit 3, 2, 1, 0: component address
 (0 - 3)
 |
Component address for correction of cross-
 |______
interference 3
Bit 7, 6, 5, 4: channel address
(1 - 12)
Bit 3, 2, 1, 0: component address
 (0 - 3)
The address for correction may not be the address of the device.
Note
The component addresses in device menu are 1 – 4 !
Read/ set
's'/'S', 39,
crc.
 |
Channel address for pressure compensation
external channel
 |_
Bit 7, 6, 5, 4: channel address
(1 - 12)
for pressure
Bit 3, 2, 1, 0 => 0
compensation
The address for compensation may not be the address of the
device.
PROFIBUS
's'/'S', 40,
crc.
profile
Bit 7: 1
 |
Bit 6,5,4,3,2,1,0 => PROFIBUS profile (0, 1, 3)
 |________
This command is only permissible with the PROFIBUS optional
board fitted.
The profile can only be set when there is no cyclical communication
.
4-36

---
Commands
Read/ set
's'/'S', 41,
crc.
 |_
Bit 7: 1
internal valves
Bit 6: 0
Bit 5: 1 => valve 6 (shed ...) open (0 => close)
Bit 4: 1 => " 5 (control air/sample gas..)"
"
Bit 3: 1 => " for
span gas
        "
"
Bit 2: 1 => " “
zero gas "
"
Bit 1: 1 => " “
combustion ai r "
"
Bit 0: 1 => " “
hydrogen
        "
"
Only one valve for zero gas, span gas and 5 can be opened at the
same time. Valve 6 can be opened only in Measure state.
Switching the valves by application flow is still active.
Valve 5 and 6 are not included for every version of FIDAMAT.
There is no valve 5 and 6 at FIDAMAT E
There is no valve 6 at FIDAMAT G.
Read
's', 42,
crc. 1, crc. 2
Relay by
 | |
Bit 7: 1
 | |___
PROFIBUS
Bit 6, 5, 4, 3, 2, 1: 0
 |
Bit 0: 1 => Relay output 8
 |
Bit 7: 1
 |________
Bit 0 - 6: Relay output 1 - 7
Relay output 1 – 8 are on the PROFIBUS optional board.
This command is only permissible with the PROFIBUS optional
board fitted and ‘Relay by PROFIBUS’ (F,33) is on.
4-37

---
 Commands
4.7
Read/ Write Values
Read values:
Command
structure
Control
 'w', number
command:
Answer:
 'w', number, val. 1 ... val. n
Example: read integration times
Control
'w', 11
command:
Answer:
'w', 11,
'10.0', 0, 9, 0,
'1.0', 0, 25, 0,
   |
'3.0', 0, 25, 0,
   |
'5.0', 0, 25, 0
   |
   | | | | |
Separator
   | | | | |_________
Devices (Table 4-1)
   | | | |___________
Separator
   | | |_____________
Values
   | |_________________
Integration times
   |_______________________
Write values:
Control
'W', number, val. 1 ... val. n
command:
Answer:
'W', number,
Example: write slope gas concentrations
Control command:
'W', 3,
'10.0', 0, '50', 0, '100.0', 0, '3000', 0
Answer:
'W', 3
Read/ write start-
'w'/'W', 1,
val. 1, val. 2, val. n
of-scale values
Start-of-scale value of highest range
  | | |
" measuring range 2
  | | |_________
" measuring range 1
  | |______________
  |___________________
≤
≤
≤
Condition: value 1
 value 2
...
 value n
With the ULTRAMAT 6 until Software version V4.2.4, value 1
is valid for all measuring ranges.
There are only value 1 to value 3 in the readonly mode available
for the C6TG.
Note
If new start-of-scale values are written, it is necessary to write the
full-scale values (”W”,2) afterwards.
4-38

---
Commands
Read/ write full-
'w'/'W', 2,
val. 1, val. 2... val. n
scale values
Full-scale value of highest range
  | | |________
"
measuring range 2
  | |_____________
"
measuring range 1
  |__________________
Condition: value 1 < value 2 <... < value n
There are only value 1 to value 3 in the readonly mode available
for the C6TG.
Read/ write slope
'w'/'W', 3,
val. 1, val. 2... val. n
gas
  | | |
Calibration gas concentration:
  | | |
concentrations
Highest measuring range
  | | |________
Measuring range 2
  | |_____________
Measuring range 1
  |__________________
The slope gas concentrations must be smaller than the full-scale
value of the associated measuring range.
There are only value 1 to value 3 in the readonly mode available
for the C6TG.
Read/ write
'w'/'W', 4,
val. 1, val. 2
linearization gas
  | |
Concentration of sensitivity gas
  | |______________
concentrations
        "
        of curve dip gas
  |__________________
This command is only valid for IR components.
The concentration of the curve dip gas must be between 30% and
70% of the full-scale value, and the concentration of the
sensitivity gas between 70% and 105% of the full-scale value.
Read/ write zero
'w'/'W', 5,
value
gas
concentration
Read/ write
'w'/'W', 6,
val. 1, val. 2, val. 3
Autocal cycle
  | | |
Carry out slope calibration following each
  | | |____
parameters
'Value 3' cycle
  | |
Cycle time (in hours)
  | |________
Time up to next Autocal (in minutes)
  |_____________
With the series 6, commands for Autocal are only valid with the
optional board fitted.
With the ULTRAMAT 23, value 3 cannot be written.
4-39

---
 Commands
Read/ write
'w'/'W', 7,
val. 1, val. 2, val. 3, ... val. 6
purge times for
  | | | |
Purge time step 6 (in min)
  | | | |______
Autocal steps 1
    "
3 "
  | | |___________
to 6
    "
2 "
  | |________________
    "
1 "
  |_____________________
Purge time inputs for step 'Signalling contact' are ignored.
With the Series 6 analyzers, commands for Autocal are only valid
with the optional board fitted.
ULTRAMAT 23 only has the Autocal steps for zero gas and
sample gas purging with the same purge time. This purge time is
entered using value 1. Further values are ignored.
If an O
 probe is connected, times less than 1 min can not be
2
entered for the ULTRAMAT 23.
Read/ write
'w'/'W', 8,
val. 7, val. 8, val. 9... val. 12
purge times for
  | | | |
Purge time step 12 (in min)
  | | | |_____
Autocal steps 7
    "
9 "
  | | |__________
to 12
    "
8 "
  | |_______________
    "
7 "
  |____________________
Purge time inputs for step 'Signalling contact' are ignored.
With the series 6 analyzers, commands for Autocal are only valid
with the optional board fitted.
Read/ write
'w'/'W', 9,
val. 1, val. 2
limit 1
  | |
Hysteresis
  | |______________
Limit
  |___________________
The hysteresis is always 0.
Read/ write
'w'/'W', 10,
val. 1, val. 2
limit 2
  | |
Hysteresis
  | |_____________
Limit
  |__________________
The hysteresis is always 0.
4-40

---
Commands
Read/ write
'w'/'W', 11,
val. 1, val. 2, val. 3, val. 4
integration times
  | | | |
Dead time (in s)
  | | | |_____
T90 outside threshold value
  | | |_________
(in s)
  | |
T90 within threshold value
  | |______________
(in s)
  |
Threshold value (in % of
  |___________________
smallest measuring range)
The dead time is always 0.
Read/ write
'w'/'W', 12,
val. 1, val. 2
autorange
  | |
Lower threshold (in % of lower measuring range)
  | |__
hysteresis
Upper threshold (in % of current measuring range)
  |_______
thresholds
The upper threshold must be higher than the lower threshold.
The function is not available for the C6TG.
Read/ write
'w'/'W', 13,
val.
 |
pump capacity
 |________
Capacity (0 - 100)
Read/ write date
'w'/'W', 14,
val. 1, val. 2, val. 3
of O
 sensor
  | | |
2
Year
(0 - 99)
  | | |______________
installation
Month
(1 - 12)
  | |___________________
Day
(1 - 31)
  |________________________
This command is only permissible with an O
 sensor connected.
2
A zero calibration (”Z”, 5) must have been executed prior to
input of date.
Read/ write time
'w'/'W', 16,
val. 1, val. 2, val. 3, val. 4, val. 5
  | | | | |
Year
 (0 - 99)
  | | | | |_____
Month
(1 - 12)
  | | | |__________
Day
(1 - 31)
  | | |______________
Hour
(0 - 23)
  | |___________________
Minute
(0 - 59)
  |________________________
Read/ write
'w'/'W', 18,
val.
 |
barometric
 |___
Read value of selected pressure sensor ('S',15), or adjust
pressure
4-41

---
 Commands
integrated pressure sensor to entered value (in hPa).
Read/ write
'w'/'W', 19,
val.
analog output
μ
By writing the measured value (in
A) at the analog output in
mode Adjust analog output, the current output is set to the
defined setpoint, and the next current value is output.
'Read analog output' shows the actual analog output current.
Read/ write LCD
'w'/'W', 20,
val.
 |
contrast
 |___
Contrast (0 - 100)
Read/ write
'w'/'W', 21,
val.
 |
frequency
 |___
Read or set the chopper frequency for the ULTRAMAT 6
and ULTRAMAT 23 or the magnetic field frequency for
the OXYMAT 6.
A Reset is executed automatically following the write command
with ULTRAMAT 23. With the ULTRAMAT 23 from Software
version V2.10.0, value is not possible to be changed.
Read/ write
'w'/'W', 22,
val.
 |
reduction value
 |___
Read reduction value for gain or set reduction value to
entered value.
Read/ write
'w'/'W', 23,
val.
 |
phase
 |___
Read phase value or set phase to entered value.
Read/ write noise
'w'/'W', 24,
val.
 |
signal
 |___
Suppress noise signals with duration of up to 'val.' (in s).
suppression
duration
4-42

---
Commands
Read/ write
'w'/'W', 25,
val. 1, val. 2, val. 3
calibration
  | | |
Number of calibrations for determination
  | | |___
tolerances
of tolerance
  | |
Calibration tolerance at zero (in % of
  | |________
smallest measuring range)
  |
Calibration tolerance of sensitivity (in %
  |_____________
of current measuring range)
The calibration tolerance of the sensitivity is not used with the
ULTRAMAT 23. The number of calibrations is not used with the
series 6 analyzers. The values have no meaning at O2-component
of ULTRAMAT 23.
Read/ write
'w'/'W', 26,
val.
  |
shock
  |__
Attenuation coefficient in the compensation circuit (in %
compensation
of the measuring circuit attenuation).
Read/ write
'w'/'W', 27,
val. 1, val. 2
parameters of
  | |
Full-scale value of external pressure sensor
  | |___
external
(in hPa)
  |
pressure sensor
Start-of-scale value
  |________
Read/ write
'w'/'W', 28,
val. 1, val. 2, val. 3
parameters of
  | | |
Full-scale value of interfering
  | | |___________
external
gas
  | |
interfering gas 1
Start-of-scale value
  | |________________
Interference
  |_____________________
With the ULTRAMAT 23, only the interference applies.
With CALOMAT 6 and FIDAMAT 6, the scale-values can only
be set when correction of residual gas influence with analog
current 1 or 2 is selected.
4-43

---
 Commands
Read/ write
'w'/'W', 29,
val. 1, val. 2, val. 3, ... val. 6
sample point
  | | | |
6th sample point time (in min)
  | | | |___
times
3rd "
     "
"
  | | |_________
2nd "
     "
"
  | |______________
1st "
      "
"
  |___________________
Inputs for sample points which are not enabled (relay set) are
ignored.
Note
The assignment of the sample point number to the sample point
time results from the number of the relay which is assigned to the
sample point. The sample point with the lowest relay number is
assigned the first sample point time, etc.
Read/ write
'w'/'W', 40,
val. 1, val. 2, val. 3
autorange lower
  | | |
Lower limit for measuring range 4
  | | |_______
limits
"
 "
" 3
  | |____________
 "
"
" 2
  |_________________
If the measured value falls below the lower limit for measuring
range x, measuring range x-1 will be selected if autorange is on.
The limit for measuring range x must be greater than or equal to
the limit of measuring range x-1.
Read/ write
'w'/'W', 41,
val. 1, val. 2, val. 3
autorange upper
  | | |
Upper limit for measuring range 3
  | | |_______
limits
"
"
" 2
  | |____________
"
"
" 1
  |_________________
If the measured value exceeds the upper limit for measuring
range x, measuring range x+1 will be selected if autorange is on.
The limit for measuring range x must be greater than or equal to
the limit of measuring range x-1. The upper limit must be greater
than the lower limit of the same measuring range.
Read/ write full-
'w'/'W', 42,
val.
 |
scale value of
 |_____
Full-scale value of linearization curve
linearization
curve
4-44

---
Commands
Read deviation
'w', 43,
val. 1, val. 2
in zero and slope
  | |
Deviation in slope calibration (not ULTRAMAT 23)
  | |__
calibration
Deviation in zero
  |______
Read deviations for last Zero calibration ('Z', 5), Adjust
component slope ('Z', 6) or Autocal ('Z', 14).
Read/ write
'w'/'W', 44,
val. 1, val. 2, val. 3
parameters of
  | | |
Full-scale value of interfering
  | | |___________
external
gas
  | |
interfering gas 2
Start-of-scale value
  | |________________
Interference
  |_____________________
With the ULTRAMAT 23, only the interference applies.
The scale-values can only be set when correction of residual gas
influence with analog current 1 or 2 is selected.
Read/ write
'w'/'W', 45,
val. 1, val. 2, val. 3
parameters of
  | | |
Full-scale value of interfering
  | | |___________
internal
gas
  | |
interfering gas 1
Start-of-scale value
  | |________________
Interference
  |_____________________
With the ULTRAMAT 23, only the interference applies.
Read/ write
'w'/'W', 46,
val. 1, val. 2, val. 3
parameters of
  | | |
Full-scale value of interfering
  | | |___________
internal
gas
  | |
interfering gas 2
Start-of-scale value
  | |________________
Interference
  |_____________________
With the ULTRAMAT 23, only the interference applies.
4-45

---
 Commands
Temperature
'w'/'W', 47,
val. 1, val. 2, val. 3, val. 4, val. 5
values of
  | | | | |
temperature 5
  | | | | |__
temperature
  | | | |________
temperature 4
compensation
  | | |______________
temperature 3
  | |____________________
temperature 2
  |_________________________
temperature 1
temperature 1 < temperature 2 < temperature 3 < temperature 4 <
temperature 5. Value 4 and value 5 are only used with
ULTRAMAT 23.
Zero point
'w'/'W', 48,
val. 1, val. 2, val. 3, val. 4, val. 5
values of
Zero point values at :
  | | | | |
temperature 5
  | | | | |__
temperature
  | | | |________
temperature 4
compensation
  | | |______________
temperature 3
  | |____________________
temperature 2
  |_________________________
temperature 1
Value 4 and value 5 are only used with ULTRAMAT 23.
If values are written, the new zero point coefficients will be
calculated automatically. Therefore it is necessary that the
temperature values ('W',47) are written before.
End point values
'w'/'W', 49,
val. 1, val. 2, val. 3
of temperature
End point values at :
  | | |
temperature 3
  | | |__________
compensation
temperature 2
  | |________________
temperature 1
  |_____________________
If values are written, the new end point coefficients will be
calculated automatically. Therefore it is necessary that the
temperature values ('W',47) and the zero point values ('W',48) are
written before.
Zero point
'w'/'W', 50,
val. 1, val. 2, val. 3, val. 4
coefficients of
  | | | |
temperature
Coefficient 4 * 1000
  | | | |____
  | | |__________
compensation
Coefficient 3 * 1000
  | |________________
Coefficient 2 * 1000
  |_____________________
Coefficient 1 * 1000
Value 4 is only used with ULTRAMAT 23.
4-46

---
Commands
End point
'w'/'W', 51,
val. 1, val. 2, val. 3
coefficients of
  | | |
  | | |__________
temperature
Coefficient 3 * 1000
  | |________________
compensation
Coefficient 2 * 1000
  |_____________________
Coefficient 1 * 1000
Value 3 is only used with OXYMAT 6.
Switching
'w'/'W', 52,
    val.
  |
temperature of
 |_____
Temperature to select the coefficients
temperature
compensation
Coefficients of
'w'/'W', 53,
val. 1, val. 2, val. 3, val. 4, val. 5
linearization
  | | | | |
Coefficient 5 * 1000
  | | | | |__
  | | | |________
Coefficient 4 * 1000
  | | |______________
Coefficient 3 * 1000
  | |____________________
Coefficient 2 * 1000
  |_________________________
Coefficient 1 * 1000
Pressure values
'w'/'W', 54,
  val. 1, val. 2, val. 3
of pressure
  | | |
high pressure
  | | |__________
compensation
medium pressure
  | |________________
  |_____________________
low pressure
Zero point
'w'/'W', 55,
 val. 1, val. 2, val. 3
values of
Zero point values at :
  | | |
high pressure
  | | |__________
pressure
medium pressure
  | |________________
compensation
  |_____________________
low pressure
If values are written, the new zero point coefficients will be
calculated automatically. Therefore it is necessary that the
pressure values ('W',54) are written before.
4-47

---
 Commands
End point values
'w'/'W', 56,
 val. 1, val. 2, val. 3
of pressure
End point values at :
  | | |
high pressure
  | | |__________
compensation
medium pressure
  | |________________
  |_____________________
low pressure
If values are written, the new end point coefficients will be
calculated automatically. Therefore it is necessary that the
pressure values ('W',54) and the zero point values ('W',55) are
written before.
Coefficients of
'w'/'W', 57,
 val. 1, val. 2, val. 3, val. 4
pressure
  | | | |
compensation
Coefficient 4 * 1000
  | | | |____
  | | |__________
Coefficient 3 * 1000
  | |________________
Coefficient 2 * 1000
  |_____________________
Coefficient 1 * 1000
Value 3 and 4 are not used with ULTRAMAT 23.
Reference of
'w'/'W', 58,
  val.
  |
pressure
 |_____
Pressure to select the coefficients
compensation
Aftercompensati
'w'/'W', 59,
 val. 1, val. 2, val. 3
on of the zero
  | | |
Compensation above the reference
  | | |__________
point
temperature (%/10°C)
  | |
Compensation below the reference
  | |________________
temperature (%/10°C)
  |
reference temperature
  |_____________________
Aftercompensati
'w'/'W', 60,
 val. 1, val. 2, val. 3
on of the
  | | |
Compensation above the reference
  | | |__________
measured value
temperature (%/10°C)
  | |
Compensation below the reference
  | |________________
temperature (%/10°C)
  |
reference temperature
  |_____________________
4-48

---
Commands
Precompen-
'w'/'W', 61,
  val.
  |
sation
 |_____
Value of precompensation
Parameters of
'w'/'W', 62,
 val. 1, val. 2, val. 3
calibration
  | | |
Maximum value
  | | |__________
Hang down value
  | |________________
Zero value
  |_____________________
Results of the calibration of hang-down and maximum
linearization.
Read/ write
'w'/'W', 63,
 val. 1, val. 2
limit 3
  | |
Hysteresis
  | |______________
Limit
  |___________________
The Hysteresis can not be entered.
Read/ write
'w'/'W', 64,
 val. 1, val. 2
limit 4
  | |
Hysteresis
  | |_____________
Limit
  |__________________
The Hysteresis can not be entered.
'w'/'W', 65,
 val. 1, val. 2
Read/ write
  | |
Autocal
Calibration tolerance of
  | |_____________
check calibration
sensitivity (in % of current
  |
tolerances
measuring range)
  |
Calibration tolerance at zero
  |__________________
(in % of smallest measuring
range)
The function is not available on the C6TG.
4-49

---
 Commands
Read/ write
'w'/'W', 66,
val. 1, val. 2, val. 3
parameters of
  | | |
Full-scale value of interfering
  | | |___________
external
gas
  | |
interfering gas 3
Start-of-scale value
  | |________________
Interference
  |_____________________
The scale-values can only be set when correction of residual gas
influence with analog current 1 or 2 is selected
Read/ write
'w'/'W', 67,
val. 1, val. 2, val. 3
parameters of
  | | |
Full-scale value of interfering
  | | |___________
external
gas
  | |
interfering gas 4
Start-of-scale value
  | |________________
Interference
  |_____________________
The scale-values can only be set when correction of residual gas
influence with analog current 1 or 2 is selected.
Read/ write
'w'/'W', 68,
  val.
  |
value of
 |_____
Value of interfering gas 1
interfering gas 1
Read/ write
'w'/'W', 69,
  val.
  |
value of
 |_____
Value of interfering gas 2
interfering gas 2
Read/ write
'w'/'W', 70,
  val.
  |
value of
 |_____
Value of interfering gas 3
interfering gas 3
Read/ write
'w'/'W', 71,
  val.
  |
value of
 |_____
Value of interfering gas 4
interfering gas 4
4-50

---
Commands
Read/ write
'w'/'W', 72,
 val. 1, val. 2, val. 3
Coefficients of
  | | |
Coefficient 3 * 1000
  | | |__________
external
  | |________________
Coefficient 2 * 1000
interfering gas 1
  |_____________________
Coefficient 1 * 1000
Read/ write
'w'/'W', 73,
 val. 1, val. 2, val. 3
Coefficients of
  | | |
Coefficient 3 * 1000
  | | |__________
external
  | |________________
Coefficient 2 * 1000
interfering gas 2
  |_____________________
Coefficient 1 * 1000
Read/ write
'w'/'W', 74,
 val. 1, val. 2, val. 3
Coefficients of
  | | |
Coefficient 3 * 1000
  | | |__________
external
  | |________________
Coefficient 2 * 1000
interfering gas 3
  |_____________________
Coefficient 1 * 1000
Read/ write
'w'/'W', 75,
 val. 1, val. 2, val. 3
Coefficients of
  | | |
Coefficient 3 * 1000
  | | |__________
external
  | |________________
Coefficient 2 * 1000
interfering gas 4
  |_____________________
Coefficient 1 * 1000
Read/ write zero-
'w'/'W', 76,
val.
 |
scale value of
 |_____
Zero-scale value of linearization curve
linearization
curve
Read/ write noise
'w'/'W', 77,
val.
 |
signal
 |___
Suppress noise signals above a level of up to 'val.' (in % of
suppression
smallest measuring range).
level
Read/ write
'w'/'W', 78,
 val. 1, val. 2
offset of
Offset of :
  | |
pressure sensor for sample gas
  | |____________
pressure
pressure sensor for H2
  |_________________
sensors
4-51

---
 Commands
Read/ write
'w'/'W', 79,
 val. 1, val. 2
parameters of
  | |
parameter 2
  | |____________
preamplifier
parameter 1
  |_________________
Heating
'w'/'W', 80,
 val. 1, val. 2, val. 3, val. 4
parameters of
  | | | |
FIDAMAT
max. deviation of oven-/
  | | | |____
  | | |
        catalyzer- heating
  | | |__________
set value of ‘flame on’
  | |________________
set value of catalyzer heating
  |_____________________
set value of oven heating
Read set values
'w', 81,
 val. 1, val. 2
of pressures
  | |
combustion air
  | |____________
hydrogen
  |_________________
Read/ write
'w'/'W', 82,
 val. 1, val. 2, val. 3
tolerances of
  | | |
difference at start of pump /
  | | |__________
pressures
change to standby
  | |
Tolerance of combustion air
  | |________________
  |_____________________
Tolerance of hydrogen
Read/ write
'w'/'W', 83,
val. 1, val. 2, val. 3, val. 4, val. 5
Drift values
  | | | | |
measuring range 4
  | | | | |__
  | | | |________
measuring range 3
  | | |______________
measuring range 2
  | |____________________
measuring range 1
  |_________________________
zero point
Drift values are the sum of calibration deviations.
With the ULTRAMAT 23, the measuring ranges 3 and 4 are not
used. The O2-component of ULTRAMAT 23 uses only
measuring range 1.
Writing drift values sets all values to 0.0. The sended values will
be ignored.
Sensor heating
'w'/'W', 84,
  val.
  |
 |_____
Set value of sensor heating
4-52

---
Commands
Read/ write
'w'/'W', 85,
 val. 1, val. 2, val. 3
offset / span of
  | | |
Span of sensor %-range
  | | |__________
sensor
  | |________________
Span of sensor vpm-range
  |_____________________
Offset of sensor
Value 3 is only valid for OXYMAT 64.
Parameter of
'w'/'W', 86,
  val.
constant cross-
  |
 |_____
Parameter of constant cross-interference
interference
Read/ write
'w'/'W', 87,
 val. 1, val. 2, val. 3, val. 4
parameters of
  | | | |
internal
Coefficient 4 * 1000
  | | | |____
  | | |__________
interfering gas 1;
Coefficient 3 * 1000
  | |________________
new
Coefficient 2 * 1000
  |_____________________
Coefficient 1 * 1000
The command is only valid on ULTRAMAT 23 with firmware
V2.14.02 or later. Coefficient 2, 3 and 4 are only valid for
component 2 or 3 and are ignored for the other.
Read/ write
'w'/'W', 88,
 val. 1, val. 2, val. 3, val. 4
parameters of
  | | | |
internal
Coefficient 4 * 1000
  | | | |____
  | | |__________
interfering gas 2;
Coefficient 3 * 1000
  | |________________
new
Coefficient 2 * 1000
  |_____________________
Coefficient 1 * 1000
The command is only valid on ULTRAMAT 23 with firmware
V2.14.02 or later. Coefficient 2, 3 and 4 are only valid for
component 2 or 3 and are ignored for the other.
4-53

---
 Commands
4.8
Read Diagnostic Values
Read values:
Command
structure
Control
'h', number
command:
Answer:
'h', number, value
Example: read temperature of measuring head (OXYMAT 6)
Control
'h', 2
command:
Answer:
'h', 2,
'75.8', 0, 40, 0
  |
| | | |
Separator
  |
| | | |_____
Device (Table 4-1)
  |
| | |_______
Separator
  |
| |_________
Value
  |
|____________
Diagnostic number
  |_________________
OXYMAT 6
Command type *
“h“, 1 Temperature of physical section
 b
        2
Temperature of measuring head
 b
        3
Detector raw signal
 k
        4
Magnetic field
 b
        5
Temperature of display
 b
        6
Calibration reserve
 k
        7
Analog input
 k
        8
Signal vector
 k
        9
Disturbance vector
 k
ULTRAMAT 6
Command type *
“h“, 1 Temperature of heating
 b
        2
Temperature of receiver
 b
        chamber
        3
Detector raw signal
 k
        4
Chopper deviation
 b
        5
Temperature of display
 b
        6
Calibration reserve
 k
        7
Analog input
 k
        8
Signal vector
 k
        9
Disturbance vector
 k
4-54

---
Commands
CALOMAT 6 / TG
Command type *
“h“, 1 Temperature of TCD
 k
        2
Not yet defined
        3
Normalized raw signal of TCD
 k
        4
Raw signal of TCD
 k
        5
Temperature of display
 b
        6
Not yet defined
        7
Analog input
 k
ULTRAMAT 23
Command type *
“h“, 1 Temperature of detectors
 k
        2
Temperature of radiator
 b
        3
Detector raw value
 k
        4
Supply raw voltage
 b
        5
Temperature of display
 b
        6
Calibration reserve
 k
        7
O
 sensor voltage
 b
2
        8
Installation voltage of the O
 b
2
   sensor
        9
Not yet defined
        10
signed measured value
 k
4-55

---
 Commands
FIDAMAT 6
Command type *
“h“, 1 Temperature of physic
 b
        2
Temperature of oven
 b
        3
Raw current
 k
        4
Temperature of catalyzer
 b
        (only at ANM-version)
        5
Temperature of display
 b
        6
Temperature of flame
 b
        7
Analog input
 k
        8
Temperature of electronic
 b
        9
Pressure of hydrogen
 b
        10
Pressure of combustion air
 b
        11
Adapterboard reference voltage
b
        12
Adapterboard negative supply
b
        13
Adapterboard positive supply
b
        14
Adapterboard high voltage
b
CALOMAT 62
Command type *
“h“, 1 Temperature of sensor
b
        2
Temperature of gas pipe (Option)
b
        3
Raw signal of sensor
b
        4
Sensor voltage
b
        5
Temperature of display
b
        6
Measured value of sensor
b
        7
Analog input
k
        8
Temperature of electronic
b
OXYMAT 64
Command type *
“h“, 1 Temperature of sensor
b
        2
Not yet defined
        3
Raw signal of sensor
b
        4
Sensor voltage
b
        5
Temperature of display
b
        6
Measured value of sensor
b
        7
Analog input
k
        8
Pressure of sample gas
b
        9
Flow of sample gas (Option)
b
*:
b = channel-related command
k = component-related command
4-56

---
Commands
4.9
Other Commands
Read measured
'k', 1,
val., crc.
value of one
  | |
Measured variable (see Tabelle 4-2)
  | |________________
component
Measured value of component
  |___________________
Tabelle 4-2
Measured variable
1.
no component
26.
C
H
6
6 (Benzol)
(only used with command "k", 2)
27.
SF
6
2.
CO
28.
CH
OH (Methanol)
3
3.
CO
29.
C
H
 OH (Ethanol)
2
2
5
4.
CH
(Methan)
30.
CH
Cl
4
2
2
5.
C
H
(Hexan)
31.
C
H
Cl
6
14
2
4
2
6.
SO
32.
CH
Cl
2
3
7.
NO
33.
C
H
O
2
4
8.
NO
34.
H
O (water vapor)
2
2
9.
CHClF
 (R22)
35.
G/l (conductivity)
2
10.
C
H
(Propan)
36.
C
3
8
11.
C
H
(Butan)
37.
S
4
10
12.
O
38.
N
2
13.
C
H
(Pentan)
39.
CF
5
12
4
14.
CnHm (THC)
40.
COCl
(phosgene)
2
15.
P (process pressure)
41.
CHF
    (R23)
3
16.
pH
42.
C
F
     (R116)
2
6
17.
T (temperature)
43.
------ (self defined
component)
18.
C
H
(Ethen)
44.
C
H
Cl
2
4
2
3
19.
C
H
(Ethin)
45.
H
  (Hydrogen)
2
2
2
20.
C
H
(Propen)
46.
Ar
3
6
21.
C
H
47.
He
4
6
22.
C
H
48.
Cl
4
8
2
23.
C
H
     (Ethan)
49. N
2
6
2
24.
NH
       (Ammoniak)
100. Help variable process pressure
3
       (only used at command ‘k’,2)
25.
N
O
2
Read measured
'k', 2,
val. 1, crc. 1, ...
val. x, crc. x
value(s) of the
  | |
channel
Measured variable (see Tabelle 4-2)
  | |_______________
Measured value of component
  |___________________
This command can be used to determine the assignment between
component and measured variable. After the components the help
variables (see Tabelle 4-2) are transmitted.
4-57

---
 Commands
Read channel
'k', 3,
crc. 1, crc. 2, crc. 3, crc. 4
functions
 | | | |
Bit 7: 1
 | | | |__
Bit 6, 5, 4 : 0
 | | |
Bit 3: 1 => CTRL at change on
 | | |
Bit 2: 1 => relay by PROFIBUS on
 | | |
Bit 1: 1 => fan on
 | | |
Bit 0: 1 => heating of catalyzer on
 | | |
Bit 7: 1
 | | |______
Bit 6: 1 => suppress negative measured
 | |
        values on the display
 | |
Bit 5: 1 => Zero calibration concerted on
 | |
Bit 4: 1 => Fault/ Maint. request/ function
 | |
        control according to NAMUR
 | |
Bit 3: 1 => heating on
 | |
Bit 2: 1 => synchronous zero calibration on
 | |
Bit 1: 1 => solenoid valve for calibration
 | |
        gas on
 | |
Bit 0: 1 => ignition on
 | |
 | |
Bit 7: 1
 | |___________
Bit 6: 1 => measured-value store on
 |
Bit 5: 1 => sample point switching on
 |
Bit 4: 1 => logbook locked
 |
Bit 3: 1 => Autocal calibration cycles on
 |
Bit 2: 1 => radiator voltage on
 |
Bit 1: 1 => pump on
 |
Bit 0: 1 => solenoid valve for zero gas on
 |
 |
Bit 7: 1
 |________________
Bit 6: 1 => measuring head heating on
Bit 5: 1 => magnetic field on
Bit 4: 1 => pressure monitor for reference /
        purge gas on
Bit 3: 1 => pressure monitor for sample gas on
Bit 2: 1 => broadcast on
Bit 1: 1 => remote on
Bit 0: 1 => maintenance switch on
4-58

---
Commands
Read component
'k', 4,
crc. 1, crc. 2, crc. 3
 | | |
functions
 | | |___
Bit 7: 1
 | |
Bit 6, 5, 4 : 0
 | |
Bit 3: 1 => Internal cross-reference 2 on
 | |
Bit 2: 1 => Internal cross-reference 1 on
 | |
Bit 1: 1 => linearization module on
 | |
Bit 0: 1 => 0
 | |
 | |______
Bit 7: 1
 |
Bit 6 : 1 => temperature compensation module on
 |
Bit 5: 1 => constant cross-interference on
 |
Bit 4: 1 => aftercompensation of the span on
 |
Bit 3: 1 => aftercompensation of the zero point on
 |
Bit 2: 1 => precompensation on
 |
Bit 1: 1 => pressure compensation on
Bit 0: 1 => temperature compensation on
 |
 |___________
Bit 7: 1
Bit 6: 1 => linearization on
Bit 5: 1 => calibrate zero before slope
Bit 4: 1 => signal violation of calibration tolerance
Bit 3: 1 => suppress negative measured values on
        the analog output
Bit 2: 1 => total calibration
Bit 1: 1 => limit monitoring on
Bit 0: 1 => autorange on
Read error state
'k', 5,
crc. 1, ... crc. x
  |
  |___________
Current errors (see Instruction Manual)
Read channel
'k', 6,
 crc., val. 1, val. 2, val. 3, val. 4, val. 5, val. 6
version
 | | | | | | |
Meas. range list
 | | | | | | |________
Firmware number
 | | | | | |_____________
Order No.
 | | | | |__________________
Revision
 | | | |______________________
Software version
 | | |__________________________
Serial No.
 | |_______________________________
Channel type
 |____________________________________
(see Table 4-3)
Values 1 to 6 may contain ASCII letters and characters (+, -, ., /,
etc.) in addition to ASCII numbers.
Table 4-3
Channel type
1
ULTRAMAT 6
2
OXYMAT 6 / OXYMAT 61
3
CALOMAT 6
4
ULTRAMAT 23
4-59

---
 Commands
5
FIDAMAT 6 E
6
occupied
7
OXYMAT 64
8
CALOMAT 62
Reset
'K', 7
linearization
coefficients
Read remaining
'k', 8,
val.
  |
time
  |__________
Remaining time for states with automatic change
upon expiry (e.g. Warm-up phase, Temperature
compensation, Zero calibration, Slope calibration...).
Read time to
'k', 9,
val.
next zero
calibration
Control command:
Read logbook
entry
'k', 10,
crc.
 |
 |___________
Entry number (last entry => number 1;
 last but one entry => number 2; etc.)
Answer:
'k', 10,
crc. 1, crc. 2, crc. 3, val. 1, val. 2, val. 3, val. 4, val. 5
 | | | | | | | |
Year
 | | | | | | | |_
Month
 | | | | | | |_____
Day
 | | | | | |___________
Hour
 | | | | |_________________
Minute
 | | | |______________________
State
 | | |____________________________
(Table 4-4)
 | |
Error No.
 | |________________________________
Entry No.
 |_____________________________________
An logbook entry without error gets error number 255.
See chapter 7.1 for error numbers.
Table 4-4
State of logbook entries
Bit 0:
1 => error appeared
0 => error disappeared
Bit 1:
1 => entry acknowledged
0 => entry not acknowledged
Bit 2, 3, 4, 5, 6:
0
Bit 7:
1
4-60

---
Commands
Acknowledge
'K', 10,
crc.
 |
logbook entry
 |___________
Entry number (last entry => number 1;
last but one entry => number 2; etc.)
Read
'k', 12,
crc. 1... crc. x
  |
maintenance
  |__________
Current maintenance requests (see Instruction
request state
Manual)
Acknowledge
'K', 13,
crc.
 |
logbook error
 |___________
Error number
Acknowledge a logbook entry with the stated error number.
See chapter 7.1 for error numbers.
If there is no logbook entry with the stated error number the
logbook remains unchanged.
Transfer data
'K', 14,
crc.
sectors
Control character
Function
1
Load user data into current working data
2
Save current working data in user data
3
Load factory settings into current working data
4
Save current working data in factory settings
5
Load basic factory data into current working data
No communication is possible while these commands are being
executed.
A Reset is executed automatically following the load commands.
With ULTRAMAT 23 the commands for the user data (control
characters 1, 2) are not possible.
Read/ enter
'k'/'K', 16,
val.
channel name
A name can be assigned to the channel to permit easy
identification.
The value may contain ASCII letters and characters (+, -, ., /,
etc.) in addition to ASCII numbers. The maximum length of the
value is 10 characters.
Read message
'k', 18,
crc. 1...
crc. x
state
  |__________
Current messages (see Instruction Manual)
4-61

---
 Commands
Read current
'k', 19,
crc.
Autocal step
Control character
Step
Comment
1
No Autocal
2
Zero gas 1
3
Zero gas 2
4
Calibration gas 1
5
Calibration gas 2
6
Calibration gas 3
7
Calibration gas 4
8
Sample gas purging
9
Intermediate sample gas mode
10
Signalling contact
Commands for Autocal are only permissible with the optional
board fitted.
Scan boot
"k", 20,
Stz.
 |___________
Bit 7: 1
program errors
Bit 6, 5, 4 : 0
Bit 3: 1 => boot program error
Bit 2: 1 => transmission error
Bit 1: 1 => data error
Bit 0: 1 => flashprom-error
Error scan only possible after termination of boot program.
Start boot
(without separator 00H)
'K', 20,
crc.
program
Control character
Step
Comment
1
Start boot program without
Only meaningful with single
checking of target address
analyzer.
2
Start boot program with
The channel must first be set to
checking of target addresses
Reset using a Reset command.
(bus mode).
This command is only accepted within one second of the Reset.
Run firmware
(without separator 00H)
'K', 21,
crc. 1, ... crc. x
update
This command is only accepted within the boot program.
Refer to Appendix for description of control characters and
sequence.
4-62

---
Commands
Clear logbook
'K', 22
All current logbook entries are cleared.
State of the
'k',23
 crc. 1, crc. 2, crc. 3, crc. 4
external
 | | | |
Bit 7: 1
 | | | |__
connections
Bit 6, 5, 4, 3, 2, 1, 0 : 0
 | | |
 | | |
External function controls
:
 | | |______
Bit 7: 1
 | |
 | |
Bit 6, 5, 4: 0
 | |
Bit 3: 1 => general 4
 | |
Bit 2: 1 => general 3
 | |
Bit 1: 1 => general 2
Bit 0: 1 => general 1
 | |
 | |
External maintenance requests :
 | |___________
Bit 7: 1
 |
 |
Bit 6: 1 => probe heating
 |
Bit 5: 1 => sample gas filter
 |
Bit 4: 1 => gas cooler
 |
Bit 3: 1 => condensate vessel
 |
Bit 2: 1 => sample pump / flow
 |
Bit 1: 1 => sample line
 |
Bit 0: 1 => other maintenance request
 |
   |____________________
External faults :
Bit 7: 1
Bit 6: 1 => probe heating
Bit 5: 1 => sample gas filter
Bit 4: 1 => gas cooler
Bit 3: 1 => condensate vessel
Bit 2: 1 => sample pump / flow
Bit 1: 1 => sample line
Bit 0: 1 => other fault
Download mode
'K'/'k', 24,
crc.
see description 'Download bei Optionskarte PROFIBUS PA/DP,
Adapterplatte FIDAMAT 6 und Physikmodul OXYMAT 64 /
CALOMAT 62'
4-63

---
 Commands
Data transfer for
'K', 25,
crc. 1...
crc. x
  |__________
Record data
download
see description 'Download bei Optionskarte PROFIBUS PA/DP,
Adapterplatte FIDAMAT 6 und Physikmodul OXYMAT 64 /
CALOMAT 62'
Read firmware
'k', 28,
val. 1, val. 2, val. 3, val. 4, val. 5, val. 6
versions
Firmware version of :
 | | | | | |
Adapterboard/Physicsmodule Boot
 | | | | | |______
Adapterboard/Physicsmodule
 | | | | |___________
 | | | |________________
Optional Board Boot
 | | |_____________________
Optional Board
 | |__________________________
Mainboard Boot
 |_______________________________
Mainboard
Not used firmware is shown with space.
Read warm-up
'k', 29,
crc.
state
Control character
Step
1
no warm-up
2
test of hydrogen
3
hydrogen not ok
4
test of combustion air
5
combustion air not ok
6
heating of oven
7
heating of chamber
8
ignition
9
start of pump
Startup state
'k'/'K', 30,
crc.
Control character
Startup state
1
Measure
2
Standby
3
Pause
Standby can not be selected with FIDAMAT 6 G.
4-64

---
Commands
Read channel
'k', 31,
crc.
variant
Channel
Variant
crc.
FIDAMAT 6
Variant E
1
"
Variant G
2
Value of crc. is 1 if channel has no variants.
Read module
'k', 32,
crc. 1, crc. 2, crc. 3, val. 1, val. 2
identification
  | | | | |
Ordering-No.
  | | | | |__
  | | | |________
Serial-No.
  | | |______________
Year of release
  | |____________________
Month of release
  |_________________________
Day of release
Module
'k'/'K', 33,
val.
user data
A name can be assigned to the sensor module to permit easy
identification.
The value may contain ASCII letters and characters (+, -, ., /,
etc.) in addition to ASCII numbers. The maximum length of the
value is 16 characters.
4-65

---
 Commands
5
Channel Modes
5.1
General
Following power-up the channels are in Warm-up mode.
When this is finished, the channels are automatically switched to
Measure mode.
5.2
Graphic Overview
Analyzer
off
Power on/
Reset
Pause
Warm-up
Measure
Standby
Calibration modes
Curve dip *
Slope
Temp.
Purge for sync. *
calibration
calibration
compensation
Lin/sens.*
zero calibr.
calibration
Flow sensor *
Zero
Synchronized *
Autocal
adjustment
calibr.
zero calibr.
Automatic change in mode
Change in mode by
(time-related functions)
command
* Only if channel includes this mode
5-66

---
Appendix
6
Examples
Read measured
Control system/PC
Analyzer
Comment
value of channel
10H, 01H,
DLE, SOH
3, component 1
30H,
Component address
D0H,
Control system address
6BH, 01H,
Command ('k', 1)
10H, 03H,
DLE, ETX
95H, C0H
Checksum
10H, 06H
DLE, ACK
10H, 01H,
DLE, SOH
D0H, 30H,
Control system address,
component address
00H,
Collective state: ready to measure
(
Table 33-1)
04H,
Mode: Measure (Table 3-2)
6BH, 01H,
Command ('k', 1)
33H, 2EH, 35H, 00H,
3.5
0BH, 00H,
% v/ v
02H, 00H,
CO
10H, 03H,
DLE, ETX
XXH, YYH
Checksum
10H, 06H
DLE, ACK
Control system/PC
Analyzer
Comment
Read error state
10H, 01H, 12H, D0H,
DLE, SOH, addresses
6BH, 05H,
Command ('k', 5)
10H, 03H, D2H, 83H
DLE, ETX, checksum
10H, 06H
DLE, ACK
10H, 01H, D0H, 12H,
DLE, SOH, addresses
05H,
Collective state: not ready to
measure, error (
Table 33-1)
01H,
Mode: Warm-up (Table 3-2)
6BH, 05H,
Command ('k', 5)
07H, 00H,
Error 7 and
1BH, 00H,
error 27 are set
10H, 03H, XXH, YYH
DLE, ETX, checksum
10H, 06H
DLE, ACK
6-1

---
Appendix
Write slope
Control system/PC
Analyzer
Comment
concentration for
10H, 01H, 13H, D0H,
DLE, SOH, addresses
channel 1,
57H, 03H,
Command ('W', 3)
component 4
32H, 30H, 30H, 2EH,
200.0
30H, 00H
10H, 03H, 6EH, FAH
DLE, ETX, checksum
10H, 06H
DLE, ACK
10H, 01H, D0H, 14H,
DLE, SOH, addresses
00H,
Collective state: ready to measure
(
Table 33-1)
04H,
Mode: Measuring (Table 3-2)
57H, 03H,
Command ('W', 3)
10H, 03H, XXH, YYH
DLE, ETX, checksum
10H, 06H
DLE, ACK
Send unknown
Control system/PC
Analyzer
Comment
command
10H, 01H, 13H, D0H,
DLE, SOH, addresses
57H, 51H,
Unknown command ('W', 81)
01H, 30H, 48H, 68H,
Any data
10H, 03H, 53H, 29H
DLE, ETX, checksum
10H, 06H
DLE, ACK
10H, 01H, D0H, 13H,
DLE, SOH, addresses
24H,
Collective state, not ready to
measure, command not accepted
(
Table 33-1)
03H,
Mode: Standby (Table 3-2)
3FH, 3FH,
'??'
10H, 03H, XXH, YYH
DLE, ETX, checksum
10H, 06H
DLE, ACK
Control system/PC
Analyzer
Comment
10H, 01H, 13H, D0H,
DLE, SOH, addresses
Send wrong
5AH, 04H,
Command ('Z', 4)
checksum
10H, 03H, XXH, YYH
DLE, ETX, checksum (wrong)
10H, 15H
DLE, NAK
6-2

---
Appendix
Channel starts automatic transmission of measured values every
Broadcasting of
500 ms.
measured values
Control system/PC
Analyzer
Comment
10H, 01H,
DLE, SOH
F0H, 30H,
Broadcast, channel address
00H,
Collective state: ready to measure
(
Table 33-1)
04H,
Mode: Measuring (Table 3-2)
6BH, 02H,
Command ('k', 2)
34H, 2EH, 31H, 00H,
Component 1: 4.1 % v/v CO
0BH, 00H, 02H, 00H,
33H, 2EH, 35H, 00H,
Component 2: 3.5 % CO
2
0AH, 00H, 03H, 00H,
31H, 30H, 31H, 33H,
Pressure: 1013 hPa process
00H, 23H, 00H, 64H,
        pressure
00H,
10H, 03H,
DLE, ETX
XXH, YYH
Checksum
There will be neither a confirmation nor an answer.
6-3

---
Appendix
7
Appendix
7.1
Error numbers for the command
      'read logbook entry' ('k',10)
error
short-
ULTRAMAT 6
OXYMAT 6
number
name
(valid from V3.0.0)
/ OXYMAT 61
(valid from V3.0.0)
1
S1
Parameter memory
Parameter memory
2
S2
Choppermotor faulty
Magnetic field supply
3
S3
Microflow sensor
Microflow sensor
4
S4
External fault
External fault
5
S5
Temp. analyzer
Temp. analyzer
6
S6
Heating defective
Heating defective
7
S7
Temperature measuring head
8
S8
Signal pressure sensor
Signal pressure sensor
9
S9
Signal too high
10
S10
24h RAM/ Flash-Check
24h RAM/ Flash-Check
11
S11
Ref. gas pressure too low
Ref. gas pressure too low
12
S12
Power supply
Power supply
13
S13
Hardware / Powerfrequency Hardware / Powerfrequency
14
S14
Measure value > full-scale value
Measure value > full-scale value
15
S15
Calibration aborted
Calibration aborted
16
S16
Sample gas flow too small
Sample gas flow too small
17
W1
Calib. diff. too large
Calib. diff. to large
18
W2
Zero calib. reserve < 20%
Sig. volt. zero to large
19
W3
Sig. volt. adj. too small
Sig. volt. adj. to small
20
W4
Set clock Fkt.58
Set clock Fkt.58
21
W5
Microflow sensor (valid until V4.2.2) Microflow sensor (valid until V4.2.2)
22
W6
Temp. sensor LCD display
Temp. sensor LCD display
23
W7
Temp. analyzer sec. > 70°C
Temp. analyzer sec. > 70°C
24
W8
Temp. meas. head > +-3°C
25
W9
External maint. request
External maint. request
26
LIM
LIM (Limit 1/ 2/ 3/ 4, channel 1)
LIM (Limit 1/ 2/ 3/ 4)
27
W10
Autocal check diff.
Autocal check diff.
28
CTRL
CTRL (Functional check)
CTRL (Functional check)
29
LIM
LIM (Limit 1/ 2/ 3/ 4, channel 2)
30
31
32
7-4

---
Appendix
error
short-
CALOMAT 6
FIDAMAT 6
number
name
(valid from V1.1.0)
(valid from V1.0.0)
1
S1
Parameter memory
Parameter memory
2
S2
Pump does not start
3
S3
Flame does not ignite
4
S4
External fault
External fault
5
S5
Temp. analyzer
Temperature oven
6
S6
Temperature catalyst
7
S7
Temperature flame
8
S8
Sample gas / H2 pressure
9
S9
Heating is switched off
10
S10
24h RAM/ Flash-Check
24h RAM/ Flash-Check
11
S11
12
S12
Power supply
Power supply
13
S13
Hardware / Powerfrequency Hardware / Powerfrequency
14
S14
Measure value > full-scale value
Measure value > full-scale value
15
S15
Calibration aborted
Calibration aborted
16
S16
17
W1
Calib. diff. too large
Calib. diff. too large
18
W2
19
W3
20
W4
Set clock Fkt.58
Set clock Fkt.58
21
W5
Sample gas / H2 pressure
22
W6
Temp. sensor LCD display
Temp. sensor LCD display
23
W7
Electronics/Physics temperature
24
W8
Flame is expired
25
W9
External maint. request
External maint. request
26
LIM
LIM (Limit 1/ 2/ 3/ 4)
LIM (Limit 1/ 2/ 3/ 4)
27
W10
Autocal check diff.
Autocal check diff.
28
CTRL
CTRL (Functional check)
CTRL (Functional check)
29
30
31
32
7-5

---
Appendix
error
short-
CALOMAT 62
OXYMAT 64
number
name
(valid from V0.1.5)
(valid from V0.1.5)
1
S1
Parameter memory
Parameter memory
2
S2
3
S3
4
S4
External fault
External fault
5
S5
Temp. analyzer
Temp. analyzer
6
S6
Heating defective (only heated versions)
7
S7
8
S8
Signal pressure sensor
9
S9
Signal too high
10
S10
24h RAM/ Flash-Check
24h RAM/ Flash-Check
11
S11
Purge gas pressure too low (Option)
12
S12
Power supply
Power supply
13
S13
Hardware / Powerfrequency Hardware / Powerfrequency
14
S14
Measure value > full-scale value
Measure value > full-scale value
15
S15
Calibration aborted
Calibration aborted
16
S16
Gas flow to small
17
W1
Calib. diff. too large
Calib. diff. too large
18
W2
19
W3
20
W4
Set clock Fkt.58
Set clock Fkt.58
21
W5
22
W6
Temp. sensor LCD display
Temp. sensor LCD display
23
W7
Heating tolerance exeeded
(only heated versions)
24
W8
25
W9
External maint. request
External maint. request
26
LIM
LIM (Limit 1/ 2/ 3/ 4)
LIM (Limit 1/ 2/ 3/ 4)
27
W10
Autocal check diff.
Autocal check diff.
28
CTRL
CTRL (Functional check)
CTRL (Functional check)
29
30
31
32
7-6

---
Appendix
error
short name
ULTRAMAT 23
number
(valid from V2.5.0)
1
S
Meas. value comp. 1 beyond tolerance
2
S
Meas. value comp. 2 beyond tolerance
3
S
Meas. value comp. 3 beyond tolerance
4
S
Meas. value O2 beyond tolerance
5
S
Mains voltage beyond tolerance
6
S
Temp. of analyzer beyond tolerance
7
S
Pressure of amb. airbeyond tolerance
8
S
Flow too low during measuring
9
S
No data for temperature compensation
10
S
Flow too low during AUTOCAL
11
S
Conc. of O2 too low during AUTOCAL
12
S
Fault at analog output
13
S
General fault of all IR-channels
14
S
Fault of adresses for IR-channels
15
S
AUTOCAL-drift beyond tolerance
16
S
Fault at EEPROM
17
S
Phase not found
18
S
19
S
20
S
IR-channel 1 not calibrated
21
S
IR-channel 2 not calibrated
22
S
IR-channel 3 not calibrated
23
S
Volt. for IR-source beyond tolerance
24
S
Supply of bridge beyond tolerance
25
S
half bridge voltage beyond tolerance
26
S
Lockin fault
27
S
External fault
28
S
29
S
30
S
31
S
Sensitivity of O2-Sensor to low
32
S
Overflow of AD-converter
33
34
35
W
AUTOCAL-drift beyond tolerance
36
W
Sensor capacity low
37
W
LCD-temperature beyond tolerance
38
W
External maint. request
39
W
40
W
The actual errorlists are in the instruction manuals.
7-7

---
Appendix
7.2
Boot Program
The channel's program memory contains a non-erasable boot
program. Using this program it is possible to download new
channel firmware by means of ELAN commands.
The boot program can be started with the boot command when
sent within one second after switching on the device or resetting
the channel.
In order to allow booting even with faulty firmware there are two
versions of the command "Start boot program" ('K', 20, 1 or 2)
1.
Boot without checking of channel address:
The boot command is accepted by all channels.
This mode is not possible in bus operation since command
answers are necessary to control the transmission.
This command version is specifically made for firmware
updates of a channel when the firmware is missing or faulty.
2.
Boot with checking of channel address:
The boot program is only accepted by the addressed channel.
In order to avoid loss of the channel address during the reset
which is required prior to booting, the reset must be executed
using the ELAN Reset command.
Booting with this command is possible during bus operation.
However, it is advisable to reduce other bus communication
(e.g. broadcast, scanning of other channels) at this time for
faster, more reliable downloading.
The command "Run firmware update" ('K', 21) is used for
downloading. The channel address is checked or not depending
on the boot command (see above).
After execution of the boot program it should be checked for an
error-free download using command ”Scan boot program errors
(”k”, 20). In case of errors restart boot program.
The firmware is stored in the PC as a file in Intel hex format
(extended). To speed up the transmission, the individual strings
(records) of this file are shortened of the start character,
checksum and end character, and converted from ASCII to binary
format.
The characters thus obtained are the data in the command.
7-8

---
Appendix
Example of data
Intel hex format
Binary format
ELAN command
transformation
('K', 21)
(write bytes 12H,
Start character:
3AH,
-
10H, 01H,
A0H, 5CH, BFH
ELAN addresses:
-
-
20H, D0H,
starting at
ELAN command:
-
-
4BH, 15H,
Data number:
30H, 34H,
04H,
04H,
address 4A80H)
Address:
34H, 41H, 38H, 30H,
4AH, 80H,
4AH, 80H,
Record type:
30H, 30H,
00H,
00H,
Data:
31H, 32H, 41H, 30H,
12H, A0H,
12H, A0H,
35H, 43H, 42H, 46H,
5CH, BFH
5CH, BFH,
Checksum:
xxH, yyH,
-
-
End character:
0DH, 0AH
-
10H, 03H,
ELAN-CRC:
-
-
XXH, YYH
1.
Switch off the analyzer, disconnect it from the bus system and
Boot procedure
connect it to the controller (PC).
without checking
of channel
2.
Switch on the analyzer and send the command 'K', 20, 1 from
address
the PC to the analyzer within the first second.
Note
The target address used is insignificant but should be within the
acceptable range (1-12) and must be retained throughout the boot
process.
3.
The analyzer is now in boot mode. The old firmware is deleted
first (takes about 20 s) and then the analyzer expects the data
transmission. While deleting no commands are accepted.
4.
Each individual firmware record must be converted in the PC
(as shown in the example) and sent to the analyzer with the
ELAN command 'K', 21. The successful transmission and the
correct programming of the record is acknowledged with DLE
(10H), ACK (06H).
If the analyzer answers with DLE (10H), NAK (15H), the
record must be transmitted again.
Note
Because duration of deleting (step 3) is undefined, it is advisable
to send the first firmware record shortly after step 3 until this
command is answered.
7-9

---
Appendix
1.
After transmission of all records the analyzer pauses for about
15 s to determine the checksum.
2.
With ELAN command ”k”, 20 it must be checked for an error-
free boot process.
3.
The new firmware can be started by switching the analyzer off
and on again or alternatively by using the ELAN Reset
command ('Z', 1).
Note
To speed up the transmission during the download and the
subsequent Reset command, there are no answering telegrams.
1.
Restart analyzer using the ELAN Reset command ('Z', 1)
Boot procedure
(takes about 20 s). No commands are answered while Reset is
with checking of
executed.
channel address
2.
Send the command 'K', 20, 2 from the PC to the analyzer
within the first second after restart.
Note
Because duration of Reset (step 1) is undefined, it is advisable to
send the command 'K',20,2 (step 2) until this command is
answered.
3.
The analyzer is now in boot mode. The old firmware is deleted
first (takes about 20 s) and then the analyzer expects the data
transmission. While deleting no commands are accepted.
4.
Each individual firmware record must be converted in the PC
(as shown in the example) and sent to the analyzer with the
ELAN command 'K', 21. The successful transmission and the
correct programming of the record is acknowledged with DLE
(10H), ACK (06H).
If the analyzer answers with DLE (10H), NAK (15H), the
record must be transmitted again.
Note
Because duration of deleting (step 3) is undefined, it is advisable
to send the first firmware record shortly after step 3 until this
command is answered.
7-10

---
Appendix
5.
After transmission of all records the analyzer pauses for about
15 s to determine the checksum.
6.
With ELAN command ”k”, 20 it must be checked for an
error-free boot process.
7.
The new firmware can be started by switching the analyzer
off and on again or alternatively by using the ELAN Reset
command ('Z', 1).
Note
To speed up the transmission during the download and the
subsequent Reset command, there are no answering telegrams.
7-11

---
Appendix
Example of boot
Control system/PC
Analyzer
Comment
program without
1.
Switch off the analyzer
checking of
2.
Switch on the analyzer and
channel address
send boot request
10H, 01H, 20H, D0H,
Start, addresses,
4BH, 14H, 01H,
command ('K', 20, 1),
10H, 03H, 89H, 51H
end, CRC
10H, 06H
3.
Analyzer accepts the request
and enters boot mode (wait
about 20 s)
4.
Send data strings
10H, 01H, 20H, D0H,
Start, addresses,
4BH, 15H,
command ('K', 21),
04H, 4AH, 80H, 00H,
data (example of
12H, A0H, 5CH, BFH,
data conversion),
10H, 03H, 26H, 6DH
end, CRC
10H, 06H
Analyzer has received data
string correctly and stored it
 => send next string
 or
10H, 15H
Error occurred
 => send string again
5.
After transmission of all
records wait approx. 10 s
6.
Send error scan after
transmission of all data strings
10H, 01H, 20H, D0H,
Start, addresses,
6BH, 14H,
command (‘k’, 20),
10H, 03H, 86H, 94H
end, CRC
10H, 06H
Analyzer has received command
correctly
10H, 01H, D0H, 20H,
Start, addresses,
04H,
State : not ready,
01H,
State : warm-up,
6BH, 14H,
command (‘k’, 20),
80H, 00H,
error state : no errors,
10H, 03H, 71H, 83H
end, CRC
7.
switch on and off or send
  command Reset
10H, 01H, 20H, D0H,
Start, addresses,
5AH, 01H,
command (‘Z’, 1),
10H, 03H, 99H, ACH
end, CRC
10H, 06H
Analyzer Reset
Analyzer starts with new firmware
7-12

---
Appendix
Example of boot
Control system/PC
Analyzer
Comment
program with
1.
Restart analyzer using Reset
checking of
command
channel address
10H, 01H, 60H, D0H,
Start, addresses,
5AH, 01H,
command ('Z', 1),
10H, 03H, 97H, 6CH
end, CRC
10H, 06H
Command received
10H, 01H, D0H, 60H,
Command answer
5AH, 01H,
10H, 03H, XXH, YYH
Analyzer carrying out a Reset
2.
Send boot request
10H, 01H, 60H, D0H,
Start, addresses,
4BH, 14H, 02H,
command ('K', 20, 2),
10H, 03H, 38H, 95H
end, CRC
10H, 06H
3.
Analyzer accepts the request
and enters boot mode (wait
about 20 s)
4.
Send data strings
10H, 01H, 60H, D0H,
Start, addresses,
4BH, 15H,
command ('K', 21),
04H, 4AH, 80H, 00H,
data (example of
12H, A0H, 5CH, BFH,
data conversion),
10H, 03H, 67H, EDH
end, CRC
10H, 06H
Analyzer has received the data
string correctly and stored it
=> send next string
 or
10H, 15H
Error occurred
=> send string again
5.
After transmission of all
records wait approx. 10 s
6. Send error scan after
     transmission of all data strings
10H, 01H, 20H, D0H,
Start, addresses,
6BH, 14H,
command (‘k’, 20),
10H, 03H, 86H, 94H
end, CRC
10H, 06H
Analyzer has received command
correctly
10H, 01H, D0H, 20H,
Start, addresses,
04H,
State : not ready,
01H,
State : warm-up,
6BH, 14H,
command (‘k’, 20),
80H, 00H,
error state : no errors,
10H, 03H, 71H, 83H
end, CRC
7. switch on and off or send
     command Reset
10H, 01H, 20H, D0H,
Start, addresses,
5AH, 01H,
command (‘Z’, 1),
10H, 03H, 99H, ACH
end, CRC
10H, 06H
Analyzer Reset
7-13

---
Appendix
Analyzer starts with new firmware
7-14